Guidance Document

Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands

Prepared for Weather Ready Pacific (WRP) Programme

Prepared by:

Shane O'Neill, Graham Elley, Andrew Harper, Molly Powers-Tora

For Weather Ready Pacific (WRP) Programme

For any information regarding this guidance document please contact:

Graham Elley

Principal Scientist

Environmental Monitoring

+64 3 343 7850

graham.elley@niwa.co.nz

New Zealand Institute for Earth Science Limited
Private Bag 14901
Kilbirnie
Wellington 6241

Phone +64 4 386 0300

| Revision | Description | Date |
| --- | --- | --- |
| Version 0.1 | Draft in preparation/in review | 24/10/2025 |
| Version 0.2 | Finalised draft for WRP stakeholders review and feedback | 30/10/2025 |
| Version 0.3 | Finalised draft for WRP stakeholders after review and feedback from Christchurch Workshop | 28/11/2025 |
| Version 0.4 | Finalised draft for WRP implementing partners after review and feedback | 29/01/2026 |
| Version 1.0 | Final version sent to client |  |
| Version 1.1 | Amendments to sections |  |
| Version 1.2 | Submitted to PICI Panel | 01/05/2026 |
| Version 2.0 | Submitted to WRP SC | 22/05/2026 |

| Quality Assurance Statement |  |  |
| --- | --- | --- |
|  | Reviewed by: |  |
|  | Formatting checked by: |  |
|  | Approved for release by: |  |

##### Executive Summary

Pacific Island Countries and Territories face some of the world’s highest exposure to severe weather, ocean, and hydrological hazards. Reliable Earth observation networks are essential for early warnings, climate monitoring, and informed decision-making. However, many Pacific National Meteorological and Hydrological Services (NMHSs) operate with fragmented, ageing, or unsustainable observation systems, constrained by limited budgets, small technical teams, and complex logistics.

The Weather Ready Pacific (WRP) Programme provides a decade-long, region-wide framework to strengthen multi-hazard early warning systems. This Guidance Document on the Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands directly supports WRP’s strategic goals by providing the technical and operational framework needed to design, procure, operate, and sustain observation networks that are fit-for-purpose, regionally interoperable, and financially sustainable.

###### The Need for a Unified Approach

Historically, observation investments across the Pacific have been project-driven, leading to incompatible systems, data silos, and unsustainable maintenance burdens. Without coordinated guidance, future investments risk duplication, stranded assets, and missed opportunities for regional collaboration.

This document addresses these risks by promoting a whole-of-system, lifecycle approach, from strategic planning and procurement through to data sharing, maintenance, and eventual renewal. It ensures that national and regional investments contribute to a cohesive, resilient Pacific observation network.

###### A Tiered, Pragmatic Framework

Recognising that not all stations serve the same purpose, this guide introduces a three-tiered system for observation infrastructure:

- Tier 1: High-accuracy reference stations for long-term monitoring.

- Tier 2: Operational stations for real-time forecasting and warnings.

- Tier 3: Cost-effective systems for targeted, high-impact applications.

This framework allows NMHSs to match infrastructure to need, optimising resources while ensuring data quality and regional compatibility.

###### Key Challenges Addressed

Through regional consultation and expert input, this guide tackles persistent barriers to sustainable observation networks:

- Fragmented procurement and incompatible technologies

- Limited long-term funding for operations and maintenance

- Gaps in technical capacity and training pathways

- Weak integration of observations into forecasting and warning systems

- Logistical and environmental challenges in remote and coastal sites

###### Strategic Outcomes

By adopting this guidance, Pacific NMHSs and partners will be better positioned to achieve:

- Interoperability: Data that flows seamlessly across national, regional, and global systems.

- Sustainability: Networks supported by realistic budgeting, spare parts strategies, and lifecycle planning.

- Capacity: A skilled, confident workforce trained by recognising capacity development as a core component of any observation system.

- Service Delivery: Observations that directly improve forecasts, warnings, and sectoral decision-making.

- Resilience: Infrastructure designed to withstand Pacific environmental conditions.

###### Core Tools and Resources

The guide is supported by a practical Procurement Toolkit, including ready-to-use templates, evaluation criteria, and a strong recommendation for “Supply & Support” contracting models to ensure ongoing technical assistance, spare parts, and calibration.

It aligns fully with WMO standards (WIGOS, GBON, RBON), the WMO Unified Data Policy, and regional initiatives such as Early Warnings for All and the Systematic Observations Financing Facility (SOFF).

###### A Call to Action

This document is a foundational instrument for regional cooperation. By working together, through shared procurement, training, maintenance, and data exchange, Pacific NMHSs can build observation networks that are greater than the sum of their parts: resilient, sustainable, and ready to meet the challenges of tomorrow.

Contents

Executive Summary	3

1	Introduction	8

1.1	Background	8

1.2	Guidance Objectives	8

1.3	Approach	9

2	End-to-End System Design and Procurement	9

2.1	Strategic Planning & Needs Assessment	11

2.2	Selection and Procurement of Equipment	11

2.3	Installation and Commissioning	11

2.4	Telemetry and ICT Infrastructure	12

2.5	Data Storage and Management	12

2.6	Integration with Forecasting and Early Warning Systems	13

2.7	End User and Service Delivery	13

2.8	Maintenance, Verification, and Calibration	13

2.9	Financing, Partnerships & Sustainability	14

2.10	Capacity Development and Training	14

3	The Tiered Framework – A Pacific Solution	16

3.1	The Three Tiers Explained	16

3.2	Mixed-Tier Stations & Tier Migration	19

3.3	Pacific-Specific Design Requirements	20

3.4	Regional Minimum Standards	20

3.5	Facilitating Effective Procurement and Capacity Building	21

4	Observing Network Variables	21

4.1	Variables	21

4.2	Variable Selection Guidance	25

5	Infrastructure & Sensor Specifications	26

5.1	Pacific Requirements	26

5.2	Tier Sensor Specification by Variable	26

5.3	Site & Installation Guidelines	33

6	Procurement ToolKit	34

6.1	Key Procurement Domains	34

6.2	Evaluation Approach	35

6.3	Pre-Procurement Considerations	35

6.4	“Supply & Support” Contracting Model	36

6.5	Procurement Templates	36

6.6	Post-Procurement Actions	37

6.7	Regional Pooled Procurement	37

7	Spare Parts, Maintenance & Lifecycle Management	38

7.1	Tiered Spare Parts Strategy	38

7.2	Spare Parts & Consumables Management	38

7.3	Maintenance Planning	39

7.4	Field Verification & Calibration	40

7.5	Traceability and Quality Assurance	40

7.6	Lifecycle Management	41

7.7	Key Recommendations	41

8	ICT, Data Management, and Telemetry Considerations	42

8.1	Telemetry Systems	42

8.2	Data Management and Sovereignty	43

8.3	Integration and Interoperability	43

8.4	Cybersecurity and Data Integrity	44

8.5	Data Sharing & Regional Collaboration	45

8.6	Capacity Development for ICT & Data Roles	45

8.7	Key Recommendations	45

9	Capacity Development	46

9.1	Training Integrated with the System Lifecycle	46

9.2	Role-Based Training Pathways & WMO Competency Alignment	46

9.3	Leveraging and Strengthening the Regional Training Ecosystem	47

9.4	Mentorship & Succession Planning	47

9.5	Mandating Training in Procurement	48

9.6	Monitoring the Impact of Capacity Development	48

9.7	Cross-Training & Regional Collaboration	49

9.8	Funding & Resourcing Capacity Development	49

9.9	Key Recommendations	49

10	Governance, Implementation & Regional Coordination	50

10.1	Roles and Responsibilities	50

10.2	Compliance Monitoring & Assurance	50

10.3	Phased Implementation Timeline	50

10.4	Governance Structure for Regional Coordination	51

11	Acknowledgements	52

12	Glossary of terms and acronyms	54

12.1	Terms and Definitions	54

12.2	Acronyms and Abbreviations	56

13	References	59

Tables

Table 1 Selection and procurement considerations for earth observations systems.	11

Table 2 Installing and commissioning considerations for earth observations systems.	12

Table 3 Telemetry and ICT infrastructure considerations for earth observations systems.	12

Table 4 Data storage and management considerations for earth observations systems.	13

Table 5 Integration with forecasting and early warning systems considerations for earth observations systems.	13

Table 6 End user and service delivery considerations for earth observations systems.	13

Table 7 Maintenance, verification and calibration considerations for earth observations systems.	14

Table 8 Financing, partnerships and sustainability considerations for earth observations systems.	15

Table 9 Capacity development and training considerations for earth observations systems.	15

Table 10 Tier 1 systems characteristics and specifications.	18

Table 11 Tier 2 systems characteristics and specifications.	19

Table 12 Tier 3 systems characteristics and specifications.	20

Table 13 Pacific-specific design requirements for all tiers to ensure resilience in tropical and coastal environments.	21

Table 14 List of key observation variables by category (Atmosphere, Marine, Freshwater & Hydrology, Land), with application areas and available technologies.	23

Table 15 General Pacific infrastructure and sensor requirements applicable to all tiers.	28

Table 16 Atmosphere (Near Surface) variable specification for each Tier (1-3). Tier 1 aligned with WMO No. 8 Annex 1.A and 1.G [1]. Tier 2 and 3 extrapolated from Tier 1 and informed by other sources, including expert judgement. Uncert. = Uncertainty; Res. = Resolution	30

Table 17 Atmosphere (Upper Air) variable specification for each Tier (1-3). Tier 1 aligned with WMO No. 8 Annex 1.A and 1.G [1]. Tier 2 and 3 extrapolated from Tier 1 and informed by other sources, including expert judgement. Uncert. = Uncertainty; Res. = Resolution	31

Table 18 Marine variable specification for each Tier (1-3). Class A aligned with WMO No. 8 Annex 1.A [1]. Tier 2 and 3 extrapolated from Tier 1 and informed by other sources, including expert judgement. Uncert. = Uncertainty; Res. = Resolution	32

Table 19 Freshwater & Hydrology variable specification for each Tier (1-3). Tier 1 is aligned with National Environmental Monitoring Standards New Zealand (NEMS) QC600 standard [9-13]. Tier 2 and 3 are extrapolated from Tier 1 based on NEMS QC500 and QC400, as well as being informed by expert judgement. Uncert. = Uncertainty; Res. = Resolution, SpC = Specific Conductivity, DO = Dissolved Oxygen	33

Table 20 Land variable specification for each Tier (1-3). Tier 1 aligned with previous tender documents in the region, as well as National Environmental Monitoring Standards New Zealand [9-13]. Tier 2 and 3 extrapolated from Tier 1 and informed by expert judgement. Uncert. = Uncertainty; Res. = Resolution, VWC = Volumetric Water Content	34

Table 21 Site selection and installation guidelines for all tiers.	35

Table 22  Key procurement domains and considerations for evaluating tenders.	36

Table 23 Example weighting scheme for procurement evaluation.	37

Table 24 Pre-procurement steps and responsibilities.	38

Table 25 Minimum requirements for "Supply & Support" contracts models.	38

Table 26 post-procurement actions and responsibilities.	39

Table 27 Tiered spare parts strategy, maintenance frequency, calibration verification frequency, and lifecycle expectations aligned with WMO Guidelines [31, 32].	41

Table 28 Three-level spare parts inventory structure (site, national, regional).	42

Table 29 Maintenance planning requirements and components.	42

Table 30 Lifecycle management phases and associated actions.	44

Table 31 Telemetry redundancy strategy options and technologies.	46

Table 32 Design specifications for resilient telemetry and power systems.	46

Table 33 Data storage options and sovereignty considerations.	47

Table 34 Archiving and metadata management requirements.	47

Table 35 Primary data integration pathways into national, regional, and global systems.	48

Table 36 Key standards for ensuring data interoperability.	48

Table 37 Cybersecurity measures for data transmission and access.	48

Table 38 Best practices for ensuring data integrity and availability.	49

Table 39 Mechanisms for regional and global data sharing.	49

Table 40 ICT and data management roles and key competencies.	49

Table 41 Training framework aligned with the system lifecycle stages.	51

Table 42 Role-based training pathways aligned with WMO competencies and tiered systems.	52

Table 43 Mentorship and succession planning strategies.	53

Table 44 Mandatory training and documentation requirements in procurement contracts.	53

Table 45  Metrics for monitoring the impact of capacity development.	54

Table 46 Cross-training and regional collaboration initiatives.	54

Table 47 Funding and resourcing sources for capacity development.	54

Table 48 Roles and responsibilities of key actors in the implementation of this guidance.	56

Table 49 Phased implementation timeline for adopting the guidance.	56

Figures

Figure 1 The End-to-End System Lifecycle. The process begins with strategic Selection and Procurement and flows through to Service Delivery, which in turn informs future procurement needs. Maintenance and Financing are not final steps but are foundational supports required for every stage of the lifecycle.	10

Figure 2 Tiered system for Earth observation infrastructure in the Pacific, categorised by purpose, accuracy, and sustainability requirements into Tier 1 (Reference), Tier 2 (Operational Forecasting), and Tier 3 (Targeted Applications). .	17

Figure 3 Governance and coordination structure for the implementation of this guidance across the Pacific region, showing roles of WRP, NMHSs, development partners, and regional technical bodies.	57

# Introduction

This Guidance Document aims to provide Pacific National Meteorological and Hydrological Services (NMHSs) and regional agencies a unified guide to help countries and partners coordinate the identification, design, procurement, installation, maintenance, and data management options for observing networks, in a way that is technically robust, regionally compatible, and sustainable.

It is a resource to support the structured enhancement and development of modern, fit for purpose hydro-meteorological data collection networks and infrastructure, and associated Infrastructure, Communications and Technology (ICT) systems, for the Pacific.

## Background

In 2021, Pacific Leaders endorsed the proposed Weather Ready Pacific (WRP) Decadal Programme of Investment, which aims to reduce the human and economic costs of severe weather, water, and ocean events by strengthening NMHSs and their partnerships with disaster management organisations.

A key component of the WRP investment strategy is to provide access to the information and resources necessary to ensure that Pacific Island Countries and Territories (PICTs) can actively participate in and benefit from global advances in early warning technologies, contribute to improved forecast Uncertainty, geographic specificity, and lead times for global and regional systems.

The ambitions of WRP are closely aligned with complementary international and regional initiatives, including:

- the UN’s Early Warnings for All (EW4ALL) initiative [2] ,

- the Systematic Observations Financing Facility (SOFF) [3] which supports Global Basic Observation Network (GBON) implementation and sustained observation networks

- the Intra-ACP Climate Services and Related Applications (CLIMSA) Program [4] supporting capacity development and climate services delivery, and

- the Climate and Oceans Support Program in the Pacific (COSPPac) [5] which has provided foundational support for ocean, tide, and climate monitoring across the region.

Collectively, these initiatives will benefit from streamlined, consistent, and regionally owned guidance that promotes interoperability, sustainability, and maximises the value of investments made at national and regional scales.

## Guidance Objectives

- Provide agreed regional guidance on the standards and principles for optimal meteorological, hydrological, and ocean observing networks and infrastructure, that is consistent with the range of station typologies that an NMHS may support to deliver its national needs.

- Support regional and national planning, procurement, and maintenance of observation and ICT systems.

- Wherever practical, promote interoperability principles, data sharing, and alignment with World Meteorological Organization’s (WMO) Unified Data Policy [6].

- Ensure sustainability by addressing total cost of ownership, regional capacity building, standardised spare parts, and lifecycle support.

- Contribute to national and regional implementation roadmaps, including SOFF and other development partner investments.

## Approach

This document will focus on observation systems for:

- Atmospheric (surface and upper-air), marine, freshwater, and land-based environmental variables.

- Supporting data management ICT and telemetry infrastructure.

- Calibration, maintenance, and lifecycle management systems.

### Methodology and Stakeholder Engagement

This Guidance Document was developed using a structured and participatory process, aimed at producing technically sound, regionally relevant, and operationally practical guidance for PICTs. The methodology acknowledges the diverse operational environments, capacities, and priorities of Pacific NMHSs, and was designed to be inclusive and responsive to their specific needs.

The development process included:

- A comprehensive review and synthesis of existing international and regional standards and frameworks, including the WMO Guide to Instruments and Methods of Observation (WMO-No. 8) [1], WMO Integrated Global Observing System (WIGOS) guidelines [7, 8], National Environmental Monitoring Standards (NEMS) [9-13] from New Zealand, and relevant regional strategies.

- Stakeholder consultations through workshops and technical discussions with NMHS directors, operational staff, regional organisations (SPREP, SPC, WMO), development partners, and technology suppliers.

- Consideration of the WRP stocktake of the current state of observing networks across the region, identifying common gaps, risks, and opportunities related to procurement, maintenance, data management, and long-term sustainability.

- Peer review by a range of technical experts from the Pacific and internationally, including meteorologists, hydrologists, ICT professionals, and procurement specialists.

This collaborative and iterative development process ensures this document is a practical and adaptive resource, designed to support Pacific NMHSs in developing resilient, interoperable, and financially sustainable observation networks that enhance early warning, climate services, and community resilience.

# End-to-End System Design and Procurement

A core principle underpinning this Guidance Document is the importance of providing the resources to enable a standardised and coordinated approach to procurement and systems integration across the Pacific region, an approach strongly advocated by WMO and other partner guidelines [14-16]. This is critical to ensure that national investments, whether made through the WRP Programme or through other bilateral and multilateral initiatives, contribute effectively to building a cohesive, interoperable, and sustainable regional observation network.

The complex and costly nature of earth observing systems in the Pacific means that poorly coordinated investments can lead to duplication, data silos, technological incompatibilities, and stranded assets. To mitigate these risks, a "whole-of-system" approach is required, one that considers the appropriate observation standards, and full lifecycle and operational context of observational infrastructure from planning to sustained operation.

Before selecting specific equipment or technologies, every Pacific NMHS must understand the complete lifecycle commitment required to operate and sustain observation systems. This section outlines the end-to-end process, from initial planning through to decommissioning, that transforms observation investments from short-term projects into long-term national assets.

The following interconnected components form the end-to-end system as shown in Figure 1.

Figure 1 The End-to-End System Lifecycle. The process begins with strategic Selection and Procurement and flows through to Service Delivery, which in turn informs future procurement needs. Maintenance and Financing are not final steps but are foundational supports required for every stage of the lifecycle.

## Strategic Planning & Needs Assessment

Before any procurement begins, NMHSs must clearly define:

- National and regional observing priorities aligned with service delivery goals

- Required variables, accuracy levels, and data timeliness

- Site-specific conditions and constraints

- Alignment with WMO frameworks (WIGOS, GBON, RBON)

- Long-term funding and capacity considerations

## Selection and Procurement of Equipment

Ensure all investments meet agreed regional standards and specifications to maximise interoperability and facilitate regional servicing and spare part supply chains. Key considerations outlined in Table 1.

Table 1 Selection and procurement considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| Fit-for-Purpose Design | Select equipment that meets multiple regional and operational needs (e.g., tropicalised, cyclone-resistant, saltwater-resistant, low power draw). |
| Regional Standards | Align with WMO Integrated Global Observing System (WIGOS) [7] and regional minimum standards to ensure interoperability, optimisation of spare parts sharing, simplified repair and maintenance and pooled systems training. |
| Total Cost of Ownership (TCO) | Include all life-cycle costs: installation, training, spare parts, maintenance, verification, calibration, telemetry, software licences, and disposal [16]. |
| Modular and Scalable Systems | Prefer modular systems that can be scaled or upgraded incrementally, reducing downtime and allowing progressive rollouts. |
| Vendor Track Record & Support | Identify vendors with proven service records in the Pacific and ensure training, warranty, and technical support are available locally or regionally. |
| Bulk procurement | Encourage the aggregation of project procurement to enable economy of scale. |

## Installation and Commissioning

Require technically sound and regionally informed installation practices that ensure data integrity from the outset. Key considerations outlined in Table 2.

Table 2 Installing and commissioning considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| Technical Guidance | Follow standard regional installation protocols to ensure data comparability, uncertainty, and system longevity. |
| Local Involvement | Engage NMHS staff and local technicians in installation to build familiarity and capacity. |
| Site Selection & Security | Choose locations with good representativeness, secure fencing, and access for maintenance. |
| Integration Planning | Ensure new systems link into national data pipelines, power sources, and ICT infrastructure. |
| Environmental Impact Assessment | Follow standard protocols to ensure negative environmental impacts are minimised. |

## Telemetry and ICT Infrastructure

Ensure reliable, secure, and regionally harmonised communication systems for data transmission, compatible with regional and global systems (e.g., WMO Information System 2.0 (WIS 2.0), WIGOS). Key considerations outlined in Table 3.

Table 3 Telemetry and ICT infrastructure considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| Data Transmission | Use reliable, redundant telemetry (e.g., GPRS, Inmarsat BGAN, Iridium, Starlink, V-SAT) suited to local conditions and bandwidth availability. |
| Power Supply | Include resilient solar systems with surge protection, battery backups, and local replacement options. |
| Data Protocols & Standards | Adhere to WIS 2.0, BUFR, and regional formats to enable real-time sharing and integration. |
| Cybersecurity & Data Integrity | Ensure systems are protected from data loss, hacking, and accidental deletion with robust authentication and backups. |

## Data Storage and Management

Prioritise systems that enable national ownership of data, while ensuring ease of regional and global data sharing in line with the WMO Unified Policy for Data Exchange [6]. Key considerations outlined in Table 4.

Table 4 Data storage and management considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| National Data Sovereignty | Data should be owned and controlled by NMHSs while being sharable regionally and globally. |
| Cloud vs On-Premises | Assess whether national servers or cloud platforms (or hybrid) offer better redundancy and affordability. |
| Archiving & Metadata | Implement data management systems that maintain complete metadata, version histories, and accessibility. |
| Data Use Cases | Systems should support end-use in forecasting, research, sectoral services (agriculture, health, Disaster Risk Reduction (DRR), and regional model assimilation. |

## Integration with Forecasting and Early Warning Systems

Design observation networks so that data flows directly into forecasting, nowcasting, and warning systems, both nationally and regionally. Key considerations outlined in Table 5.

Table 5 Integration with forecasting and early warning systems considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| Real-Time Data Ingestion | Ensure data can be ingested automatically into forecast models, warning platforms, and visualisation tools with minimal latency. |
| Compatibility with Models | Select sensors and data types that align with forecast model inputs (e.g., WRF, hydrological models) and ensure required variables are collected. |
| Forecast System Readiness | Train forecasting teams to use new data sources and integrate them into workflows, ensuring data translates into improved services. |
| Regional Data Sharing | Ensure national data supports regional forecasting efforts (e.g., Regional Specialised Meteorological Centres (RSMCs), Pacific Ocean forecast systems) via WIS 2.0-compliant sharing. |

## End User and Service Delivery

Ensure that observations ultimately serve users, including disaster managers, planners, mariners, farmers, and others, with timely, accessible, and usable information products. Key considerations outlined in Table 6.

Table 6 End user and service delivery considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| User-Centric Design | Systems should be designed with forecast and warning service needs in mind, not just data collection. Work with key stakeholders during planning stages to ensure infrastructure and data outputs meet real-world decision-making needs. |
| Feedback Loops | Enable user input (e.g., disaster managers, ocean users, health sectors) into system design, performance review, and improvements. |
| Communication Tools | Invest in the capacity to convert observations into accessible, timely, and trusted information products. |

## Maintenance, Verification, and Calibration

Use the TCO approach to plan for long-term maintenance, including spare parts, calibration services, training, and ICT upgrades. Key considerations outlined in Table 7.

Table 7 Maintenance, verification and calibration considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| Maintenance Planning | Develop national maintenance schedules and Standard Operating Procedures (SOPs), supported by trained staff and stocked spares. |
| Field Verification | Process of evaluating and confirming the accuracy, reliability and operational performance of measuring instruments at their operational site against pre-defined requirements by comparing their measurements to reference or portable standards. This verification is carried out under actual environmental conditions to ensure that instruments continue to provide reliable, consistent, and accurate data after installation. |
| Calibration Infrastructure | Access to regional or national calibration laboratories, including for pressure, radiation, temperature, field verification (travelling) standards, and radar. |
| Inventory Management | Maintain digital records of assets, serial numbers, warranties, spares, maintenance history, and fault logs. |
| Environmental Protection | Regular cleaning, corrosion control, cyclone-proofing, and pest protection to extend asset life. |
| Asset Information System | Up to date information on equipment and maintenance status is maintained and available via an asset information system |

## Financing, Partnerships & Sustainability

Sustainable observation networks require long-term financial planning, coordinated partnerships, and alignment with regional standards. Investments should account for total lifecycle costs – procurement, operations, maintenance, and calibration. Leveraging regional procurement, shared services, and aligned donor support enhances cost-efficiency and system compatibility. Strong Monitoring and Evaluation (M&E) systems ensure continued relevance and performance of infrastructure. Key considerations outlined in Table 8.

Table 8 Financing, partnerships and sustainability considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| National Budgeting | Advocate for dedicated Operating and Maintenance (O&M) funding lines in national budgets as part of climate resilience or DRR investment. |
| Regional Procurement | Use regional platforms (e.g., SPREP, PIFS, WMO) for joint or bulk procurement and cost-sharing. |
| Donor Coordination | Harmonise donor support to avoid duplication, conflicting technologies, or unsustainable systems. |
| Monitoring & Evaluation | Regularly evaluate station uptime, data quality, capacity metrics, and end-user` satisfaction to inform investment updates. |

## Capacity Development and Training

A sustainable observation network is dependent on a skilled, confident, and well-supported workforce [14, 15]. Capacity development in the Pacific must be strategic, continuous, and tailored to the specific roles within the end-to-end system, from technicians and forecasters to data managers and ICT specialists. A focus on standardised regional training, knowledge sharing, and succession planning is critical to mitigate staff turnover and build enduring institutional capability. Key considerations outlined in Table 9.

Table 9 Capacity development and training considerations for earth observations systems.

| Consideration | Detail |
| --- | --- |
| Needs-Based Training Strategy | Develop and regularly update national and regional training plans based on systematic skills gap analyses, aligned with WMO competency frameworks [15] and the specific technology tiers (1, 2, 3) in use. Training must be specific: - Technicians: Equipment-specific maintenance and repair, field verification, grounds maintenance, basic telemetry troubleshooting. - Data Managers: Data QC procedures, metadata standards, use of data management systems (e.g., CliDE). - Forecasters: Interpretation and application of new data streams in forecasting and warning workflows. - Upper-Air Specialists: Radiosonde system operation, maintenance, and data handling. |
| Leverage Regional Centres | Prioritise and utilise WMO Regional Training Centres (RTCs) like RTC Nadi and RTC Melbourne, and partner with regional institutions (e.g., USP, SPC) for foundational and advanced training, fostering a cohesive Pacific expertise base. |
| Hands-On and Localised Training | Ensure training curricula include substantial practical, hands-on components relevant to Pacific conditions (e.g., cyclone-proofing, corrosion control, vegetation management, troubleshooting in remote settings). |
| Integrate Training with Procurement | Mandate that equipment suppliers provide comprehensive initial training and detailed documentation as a core requirement of all procurement contracts, especially under "Supply & Support" models. |
| Continuous Professional Development | Establish programs for ongoing learning beyond initial installation, including workshops on new technologies, data quality control, maintenance Standard Operating Procedures, and the use of observations in service delivery (e.g., for early warnings). |

# The Tiered Framework – A Pacific Solution

With a clear understanding of the observation lifecycle, NMHSs can make informed decisions about the appropriate level of investment and commitment for different observation needs. To ensure practicality and cost-effectiveness, this Guidance Document adopts a three-tiered approach to Pacific Earth observation system standards (shown in Figure 2).

This framework emerged from stakeholder feedback highlighting the need for flexible, purpose-driven standards that match varying national capacities, priorities, and resources. The tiered system is not about creating a hierarchy of importance, but rather a hierarchy of purpose and sustainability. It is a pragmatic and strategic tool that enables Pacific NMHSs to build resilient, fit-for-purpose, and financially sustainable observation networks that directly serve national priorities and contribute effectively to regional resilience. The tiered system ensures that limited resources are strategically allocated. Not every station needs to meet the highest international standards for climate benchmark monitoring. This approach allows NMHSs to design a cost-effective network tailored to specific applications.

Tier 1

Summary: Stations meeting the highest accuracy standards for long-term monitoring, serving as national. These stations require the highest specification instruments and most rigorous maintenance schedules.

Purpose: To establish and maintain definitive, long-term climate records across the atmosphere, ocean, and land; and to contribute to global climate assessment and forecasting systems.

Tier 2

Summary: Stations designed for operational weather forecasting and warning. They meet all WMO Regional Basic Observing Network (RBON) requirements and provide high-quality, reliable data for models and nowcasting.

Purpose: To provide real-time, reliable data for daily forecasting and severe weather warnings, including cyclones, storm surges, floods, and high seas.

Tier 3

Summary: Stations deployed for specific, high-impact applications (e.g., rainfall in flash-flood prone catchments, soil temperature and moisture monitoring for agriculture, etc.). The focus is on robustness, reliability, and cost-effectiveness for a targeted purpose, potentially relaxed accuracy specifications for non-critical parameters.

Purpose: To address specific, local risks or sectoral needs with a practical, cost-effective solution.

Figure 2 Tiered system for Earth observation infrastructure in the Pacific, categorised by purpose, accuracy, and sustainability requirements into Tier 1 (Reference), Tier 2 (Operational Forecasting), and Tier 3 (Targeted Applications).

## The Three Tiers Explained

###### TIER 1: REFERENCE STATIONS

Purpose: To establish and maintain definitive, long-term records across the atmosphere, ocean, and land; and to contribute to global climate assessment and forecasting systems. Key characteristics and specification outlined in Table 10.

Important Note: This document uses the terms Tier 1, 2, and 3 to define operational categories of observing stations. In regard to atmospheric observations, this is distinct from the WMO Measurement Quality Classifications (A, B, C, D) defined in WMO-No. 8  ADDIN EN.CITE <EndNote><Cite><Author>WMO</Author><Year>2025</Year><RecNum>1</RecNum><DisplayText>[1]</DisplayText><record><rec-number>1</rec-number><foreign-keys><key app="EN" db-id="5drtpt0wbe95whe50fb5rs0dvsrxsvdrz29t" timestamp="1761767552">1</key></foreign-keys><ref-type name="Generic">13</ref-type><contributors><authors><author>WMO</author></authors></contributors><titles><title>Guide to Instruments and Methods of Observation (WMO-No. 8)</title></titles><dates><year>2025</year></dates><urls><related-urls><url>https://library.wmo.int/idurl/4/41650 </url></related-urls></urls></record></Cite></EndNote>[1], which specify the realised uncertainty of individual measurements, accounting for instrument, calibration, environmental, and maintenance factors. A station in this Guidance Document’s operational Tier 1 will typically contain sensors aiming for WMO Measurement Quality Class A or B, but the "Tier" encompasses the entire system's purpose, design, and support model.

Table 10 Tier 1 systems characteristics and specifications.

| Characteristic | Tier 1 Specification |
| --- | --- |
| Primary Purpose | Long-term climate monitoring and trend analysis; Contribution to GBON; Climate research and policy baseline |
| Key Applications | - Climate change detection - Global climate assessments (IPCC) - Sea-level rise reference monitoring - National climate policy development - Calibration reference for other stations |
| Measurement Accuracy | Highest technical specifications |
| Data Completeness | >98% data recovery; Gaps must be documented and justified |
| Redundancy | Multiple sensors for critical variables; Backup power and communications |
| Maintenance Schedule | Monthly visual checks; Quarterly detailed inspections; Annual full calibration |
| Calibration Requirements | Annual calibration against national/international standards; Traceable to BIPM |
| Spare Parts Strategy | Full redundancy at site; National inventory of all critical components |
| Technical Skill Requirements | Advanced training in metrology; Specialised calibration skills; Data quality expertise |
| Site Requirements | Long-term stable location (>20 years); Minimal environmental interference; Comprehensive site metadata; Secure from vandalism; Accessible for maintenance |
| Data Transmission | Real-time via redundant links (cellular + satellite); Complete daily datasets transmitted |
| Metadata Requirements | Full WIGOS metadata standard; Complete maintenance and calibration history |
| Lifecycle Expectation | 10-12 years with mid-life refurbishment |
| Typical Sites | National climate reference stations; GBON stations; Primary tide gauges; Upper-air reference stations |

###### TIER 2: OPERATIONAL FORECASTING STATIONS

Purpose: To provide real-time, reliable data for daily forecasting and warnings, including tropical cyclones, storm surges, floods, and high-impact weather events. Key characteristics and specification outlined in Table 12.

Table 11 Tier 2 systems characteristics and specifications.

| Characteristic | Tier 2 Specification |
| --- | --- |
| Primary Purpose | Real-time severe weather warnings; Operational forecasting input; Aviation and maritime safety |
| Key Applications | - Tropical cyclone tracking - Flash flood warnings - Aviation weather services - Marine forecasts and warnings - Agricultural advisories |
| Measurement Accuracy | Meets WMO RBON RA IV requirements |
| Data Completeness | >95% data recovery; Critical during severe weather events |
| Redundancy | Backup power essential; Alternative communications pathway recommended |
| Maintenance Schedule | Monthly visual checks; Bi-annual detailed inspections |
| Calibration Requirements | Bi-annual verification; Annual calibration for critical sensors |
| Spare Parts Strategy | Essential spares at national level; Regional pooling for expensive items |
| Technical Skill Requirements | Standard meteorological technician training; Basic calibration skills; Fault diagnosis capability |
| Site Requirements | Representative of area; Secure from vandalism; Accessible for maintenance; Comprehensive site metadata |
| Data Transmission | Real-time via primary link; Daily synoptic transmissions; ≤15-minute latency for warnings |
| Metadata Requirements | Complete WIGOS operational metadata; Maintenance records |
| Lifecycle Expectation | 8-10 years with regular maintenance |
| Typical Sites | Airport meteorological stations; Tsunami warning stations; Flood forecasting points; Cyclone monitoring locations |

###### TIER 3: TARGETED APPLICATION STATIONS

Purpose: To address specific, local risks or sectoral needs with a practical, cost-effective solution where the highest accuracy specifications may be relaxed for non-critical parameters. Key characteristics and specifications outlined in Table 12.

Table 12 Tier 3 systems characteristics and specifications.

| Characteristic | Tier 3 Specification |
| --- | --- |
| Primary Purpose | Specific local risk monitoring; Community-based applications; Sectoral decision support |
| Key Applications | - Community rainfall monitoring - Agricultural soil moisture - Small catchment flood alerts - Localised wind for renewable energy - Water resource management |
| Measurement Accuracy | Fit-for-purpose accuracy |
| Data Completeness | Gaps acceptable if documented |
| Redundancy | Minimal; Single power and communication systems |
| Maintenance Schedule | Quarterly visual checks; Annual inspections; Reactive maintenance approach |
| Calibration Requirements | Factory calibration with periodic verification (2-3 years); "As-failed" recalibration |
| Spare Parts Strategy | Minimal spares; Vendor-supported replacement; Regional pooling for common items |
| Technical Skill Requirements | Basic maintenance training; Community-based technicians acceptable |
| Site Requirements | Practical location for application; Community-managed sites acceptable |
| Data Transmission | Daily or weekly transmissions acceptable; Cellular or manual download |
| Metadata Requirements | Minimum metadata for interpretation; Site and sensor documentation |
| Lifecycle Expectation | 5-7 years; Lower-cost replacement strategy |
| Typical Sites | Community rain gauges; School weather stations; Agricultural monitoring; Small river level gauges; Coastal erosion points |

## Mixed-Tier Stations & Tier Migration

###### Mixed-Tier Stations

A single site may include instruments from different tiers (e.g., Tier 1 climate sensors alongside Tier 3 application sensors). Each instrument stream should be clearly documented with its tier classification.

###### Tier Migration

Stations may change tiers over time due to:

- Changing national priorities

- Improved technical capacity

- Budget constraints

- Evolving risks

Migration plans should include:

- Assessment of current vs. desired tier specifications

- Budget implications for upgraded maintenance and calibration

- Training requirements for staff

- Timeline for transition with minimal data disruption

## Pacific-Specific Design Requirements

All tiers must include Pacific-specific adaptations to ensure resilience in the region's challenging environments. Details outlined in Table 13.

Table 13 Pacific-specific design requirements for all tiers to ensure resilience in tropical and coastal environments.

| Requirement | Specification |
| --- | --- |
| Cyclone Resistance | Structures and mounts designed to withstand Category 4+ cyclonic winds |
| Salt Spray Protection | Corrosion-resistant materials (marine-grade stainless steel, powder-coated aluminium) |
| UV Stability | UV-stable enclosures and sensor materials |
| Waterproofing | Minimum IP65 rating for all outdoor components; IP68 for coastal/marine installations |
| Solar Power Compatibility | Designed for off-grid solar operation with battery backup |

## Regional Minimum Standards

While the tiers have different performance levels, they all operate within a single, unified regional framework. This promotes interoperability by ensuring that data from a Tier 3 station is still shared in standard formats and can be understood and used alongside data from a Tier 1 station. It prevents technological fragmentation and supports the goal of a cohesive system for the Pacific.

All tiers must meet minimum regional standards for:

- Standardised data formats and transmission protocols (WIS 2.0)

- Complete WIGOS metadata required

- CliDE, or other NMHS preferred database, compatible systems to ensure secure and sovereign data collection

- WRP regional observation standards outlined in this Guidance Document

## Facilitating Effective Procurement and Capacity Building

The tier specifications provide a clear basis for procurement. Partners and NMHSs can issue tenders with confidence that all bids will be evaluated against the same technical criteria.

Furthermore, training and capacity development can be tailored to tier requirements:

- Tier 1 Technicians: Advanced calibration, meteorology, data quality assurance

- Tier 2 Technicians: Routine maintenance, troubleshooting, basic calibration

- Tier 3 Technicians/Community Observers: Basic operation, visual checks, data collection

This allows for more targeted and efficient workforce development, ensuring that limited training resources are allocated where they are most needed.

# Observing Network Variables

This section provides a prioritised list of variables essential to support PICTs’ national mandates. These align with:

- WMO’s GBON guidelines [7]

- Regional Basic Observing Network (RBON) guidelines for Regional Association (RA) V [17] which encompasses the South-West Pacific

- The Pacific Island Communications, Infrastructure (PICI) Panel’s [18] focus areas

- Regional consultation outcomes from the WRP programme

The following tables categorise variables by domain, indicating their key applications and available technologies.

## Variables

Table 14 List of key observation variables by category (Atmosphere, Marine, Freshwater & Hydrology, Land), with application areas and available technologies.

| Category | Variables | Key Application Areas | Available technologies (non-exhaustive) |
| --- | --- | --- | --- |
| Atmosphere (Near Surface) | Air Temperature (Instantaneous) | Real-time monitoring for forecasting, agriculture, health (heat stress), aviation, and climate monitoring | Integrated AWS sensors (Temperature probe/Digital thermometer) Manual Observers (alcohol thermometers or digital with display) |
|  | Daily Maximum Air Temperature | Heatwave and health warnings, climate records, agricultural stress indices, fire weather indices | Integrated AWS sensors (Temperature probe/Digital thermometer with max temperature logging capability) Manual Observers (alcohol thermometers or digital with display) |
|  | Daily Minimum Air Temperature | Frost monitoring, agricultural impacts (crop damage), climate records, public health (cold stress) | Integrated AWS sensors (Temperature probe/Digital thermometer with min temperature logging capability) Manual Observers (alcohol thermometers or electronic handhelds used instead) |
|  | Air relative humidity | Weather warnings, health, fire weather | Integrated AWS sensors (Humidity Probe/Digital hygrometer/Wet bulb temperature probe) Manual Observers (dry and wet bulb thermometer or digital with display) Radiosondes |
|  | Temperature dew point | Forecasting, Aviation | Integrated AWS sensors (Dew point probe/Digital hygrometer/Wet bulb temperature probe) Manual Observers (dry and wet bulb thermometer or digital with display) |
|  | Air pressure | Forecasting, storm surge modelling | Integrated AWS sensors (digital barometers) Standalone Barometers/Barographs |
|  | Wind speed and direction | Cyclone monitoring, aviation, maritime safety | Integrated AWS sensors (Anemometer, cup counter, propeller, Ultrasonic anemometer) Radiosondes |
|  | Wind gust |  |  |
|  | Accumulated precipitation | Hydrology, flood early warnings, agriculture | AWS (Tipping Bucket Rain Gauge TBRG) Manual Observers (manual rain gauge) Pluviograph Standalone TBRG Radars |
|  | Precipitation intensity |  |  |
|  | Solar radiation & sunshine duration | Energy generation, agriculture productivity, modelling climate, public health assessment, urban planning | Sunshine recorders (e.g. Campbell-Stokes) Silicon photodiode sensors, All-sky cameras Integrated AWS radiation sensors Satellite remote sensing |
| Atmosphere (Upper Air) | Atmospheric Temperature (Profile) | Forecast modelling, aviation, cyclone prediction | Radiosondes Pilot balloons Aircraft observations (Doppler) Weather Radar Wind profiler radar GNSS-Radio Occultation (RO) Surface-based remote sensing stations such as Microwave Radiometers, Raman Lidar, DIAL |
|  | Atmospheric Pressure (Profile) | Synoptic analysis, model assimilation |  |
|  | Relative humidity (Profile) | Cloud formation, storm development |  |
|  | Wind/wind shear | Jet streams, vertical wind shear analysis |  |
| Marine | Sea surface temperature (SST) | Coral bleaching monitoring, ocean heat content tracking, cyclone forecasting, fisheries management | In situ thermistors (Drifting/Moored Wave Buoys, Argo Floats) Manual Measurements Satellite infrared sensors |
|  | Coastal sea level (tide) including: - storm surge, - coastal sea level, - coastal inundation, and - sea surface anomaly | Coastal flood forecasting, infrastructure planning & zoning, tide predictions for navigation & fishing, sea level rise monitoring | Water (Tidal) Level Gauges (radar, bottom pressure) High-frequency (HF) near-shore water level sensors Satellite scatterometers |
|  | Sea state including: - significant wave height, - dominant (or peak) wave period, - mean wave period, - dominant (or peak) wave direction and - mean direction | Marine safety and navigation, port operations, coastal erosion forecasting, wave energy potential | HF radars Pressure sensors Acoustic Doppler Current Profiler (ACP) Wave Buoys Manual Observations |
| Freshwater & Hydrology | River Height / Discharge | Flood early warning systems, water resource allocation, dam and reservoir operations, hydrological modelling | Pressure or radar water level sensors Acoustic Doppler sensors Staff gauges (manual) |
|  | Groundwater Levels | Drought monitoring, sustainable water extraction, climate change impact assessment, agricultural planning | Pressure transducers in boreholes Datalogger systems for continuous monitoring |
|  | Water Quality | Drinking water safety, pollution control, public health monitoring, ecosystem health assessments | In situ multivariable sondes Handheld field meters Lab analysis kits |
| Land | Soil moisture | Agricultural productivity & drought monitoring, irrigation scheduling, landslide risk assessment, flash flood warning | Integrated AWS sensors (capacitance sensors) Remote sensing |
|  | Ground (earth) temperature | Agricultural productivity & drought monitoring, soil & ecosystem health, climate change & carbon monitoring, environmental monitoring, landslide risk assessment | Integrated AWS sensors (thermistors) Data loggers with temperature probes Remote sensing |
|  | Land Use / Cover | Deforestation tracking, urban expansion monitoring, disaster impact assessments, emissions reporting & carbon stock estimates, hydrological modelling | Satellite imagery (e.g., Sentinel-2, Landsat) Aerial surveys Field surveys |

## Variable Selection Guidance

When planning an observing network, consider:

- Service-Driven Selection

- Which variables are needed for specific services (e.g., flood warnings → rainfall, river level)?

- Which user groups depend on which variables?

- Tier-Appropriate Instrumentation

- Tier 1: Highest accuracy, redundancy, regular calibration

- Tier 2: Operational reliability, real-time transmission

- Tier 3: Fit-for-purpose, cost-effective, low maintenance

- Integration Potential

- Can one sensor measure multiple variables (e.g., multi-parameter sonde)?

- How will data flow into national and regional systems?

- Sustainability

- What are the long-term costs for calibration, maintenance, and replacement?

- Is there regional capacity to support this variable's measurement?

# Infrastructure & Sensor Specifications

This section outlines the recommended specifications for observation infrastructure and sensors across marine, freshwater, land, and atmospheric domains for the Pacific.

All infrastructure investments should align with regional and WMO standards, where possible, ensure compatibility with national forecasting and early warning systems, and follow a whole-of-system approach that supports sustainability, capacity development, and data-sharing.

## Pacific Requirements

All observing equipment and systems (all tiers) should meet the following minimum general requirements outlined in Table 15.

Table 15 General Pacific infrastructure and sensor requirements applicable to all tiers.

| Requirement Category | Specification |
| --- | --- |
| Pacific-Resilient Design | Equipment must be corrosion-resistant, UV-stable, cyclone-resilient, and waterproof to at least IP65 standards. Coastal/marine deployments must withstand salt spray and high humidity. |
| Power Supply | Solar-powered with minimum 14 days battery backup. Must include surge protection, voltage regulation, and local replacement capacity. |
| Telemetry | Reliable and redundant communications (e.g., GPRS, 3G/4G, Inmarsat BGAN, Iridium, VSAT, Starlink) compatible with national ICT infrastructure. Must support data transmission to CliDE or equivalent NMHS databases. |
| Data Storage | On-device storage for ≥1 year of data. Support for open data formats (CSV, JSON, WMO BUFR/netCDF). |
| Calibration Readiness | Designed for regular calibration cycles with accessible ports/connectors. Support for field verification using portable standards. |
| Vendor Support | Minimum 3-year warranty. Access to spare parts, local/regional technical support, remote diagnostics. Vendor must have ≥3 years’ experience operating in PICTs with ≥5 successful tropical/marine installations. Regional office or dedicated Pacific support team preferred. |

## Tier Sensor Specification by Variable

The Tier 1, 2, and 3 values in Table 16 to Table 20 are derived from and aligned with primarily with:

- WMO Guidance [1, 19] for atmosphere near surface and upper air

- National Environmental Monitoring Standards (NEMS) New Zealand [9-13] for freshwater and hydrology

- Technical specifications developed under the Intra-ACP CLIMSA Programme which were widely circulated and reviewed by regional experts from WMO, SPREP, BoM, Earth Sciences New Zealand, and Pacific NMHSs. Covers atmosphere near surface and land.

- Manufacturing specifications for equipment typically used in the systems [20-30].

Practical experience with instrument performance in Pacific environments (high humidity, salt spray, intense UV, tropical rainfall).

Where sensor specifications were not available from WMO guidance, defined in previous regional projects, or from other primary sources, regional expert opinion has been relied upon.

Important Note on Specifications and Uncertainty: The uncertainties specified in Tables 2-6 represent sensor performance under ideal test conditions. The realised system uncertainty will be higher due to environmental exposure, installation effects, maintenance schedules, and calibration drift. These specifications are designed as a procurement tool to select instruments capable of meeting the intended operational tier's purpose and data quality goals, as defined by the WMO Measurement Quality Classifications.

Use WMO Measurement Quality Classifications (A, B, C, D) defined in WMO-No. 8  ADDIN EN.CITE <EndNote><Cite><Author>WMO</Author><Year>2025</Year><RecNum>1</RecNum><DisplayText>[1]</DisplayText><record><rec-number>1</rec-number><foreign-keys><key app="EN" db-id="5drtpt0wbe95whe50fb5rs0dvsrxsvdrz29t" timestamp="1761767552">1</key></foreign-keys><ref-type name="Generic">13</ref-type><contributors><authors><author>WMO</author></authors></contributors><titles><title>Guide to Instruments and Methods of Observation (WMO-No. 8)</title></titles><dates><year>2025</year></dates><urls><related-urls><url>https://library.wmo.int/idurl/4/41650 </url></related-urls></urls></record></Cite></EndNote>[1] to account fully for the realised uncertainty of individual measurements, accounting for instrument, calibration, environmental, and maintenance factors.

### Atmosphere (Near Surface)

Table 16 Atmosphere (Near Surface) variable specification for each Tier (1-3). Tier 1 aligned with WMO No. 8 Annex 1.A and 1.G [1]. Tier 2 and 3 extrapolated from Tier 1 and informed by other sources, including expert judgement. Uncert. = Uncertainty; Res. = Resolution

| Category | Variable | Tier 1 | Tier 2 | Tier 3 | Comments |
| --- | --- | --- | --- | --- | --- |
| Atmosphere (Near Surface) | Air temperature (Instantaneous) | Uncert.: ±0.2 °C Res.: 0.1 °C Range: -40 to +60°C | Uncert.: ±0.6 °C Res.: 0.2 °C Range: -20 to +60°C | Uncert.: ±1.0 °C Res.: 0.5 °C Range: -20 to +60°C | Sensor to be installed outdoors in a thermometer screen with sensor at a height of 1.25 m minimum above ground level. |
|  | Daily maximum / minimum air temperature | Uncert.: ±0.1 °C Res.: 0.1 °C | Uncert.: ±0.2 °C Res.: 0.1 °C | Uncert.: ±0.5 °C Res.: 0.1 °C | Daily max derived from continuous records or digital thermometers |
|  | Air relative humidity | Uncert.: ±2% Res.: 1% Range: 0–100% | Uncert.: ±5% Res.: 2% Range: 0–100% | Uncert.: ±10% Res.: 5% Range: 0–100% | Sensors must be stable in high humidity and salt air. Co-located with temperature sensor in same shield. |
|  | Air pressure | Uncert.: ±0.2 hPa Res.: 0.1 hPa Range: 500–1080 hPa | Uncert.: ±1.0 hPa Res.: 0.2 hPa Range: 500–1080 hPa | Uncert.: ±2.0 hPa Res.: 0.5 hPa Range: 500–1080 hPa | Station level and mean sea level to be reported. |
|  | Wind speed | Uncert.: ±0.5 m/s for wind speed ≤5 m/s and ±10% for a wind speed >5 m/s Res.: 0.5 m/s or better Range: 0–75 m/s | Uncert.: ±1 m/s for wind speed ≤5 m/s and ±15% for a wind speed >5 m/s Res.: 0.5 m/s or better Range: 0–75 m/s | Uncert.: ±2.5 m/s for wind speed ≤5 m/s and ±20% for a wind speed >5 m/s Res.: 1.0 m/s or better Range: 0–60 m/s | Cup or ultrasonic anemometers. Standard exposure is at a height of 10 m above flat ground. Distance from obstructions-to-free-flow-of-air by a minimum of 10 (preferably 30) times the height of the obstruction. |
|  | Wind gust | Uncert.: same as wind speed Res.: same as wind speed Range: 0.1–75 m/s Sampling interval: ≤3 secs | Uncert.: same as wind speed Res.: same as wind speed Range: 0.1–120 m/s Sampling interval: ≤3 secs | Uncert.: same as wind speed Res.: same as wind speed Range: 0.1–100 m/s Sampling interval: ≤5 secs | Gust is defined as a 3-second average. |
|  | Wind direction | Uncert.: ±5° Res.: 1° Range: 0–360° | Uncert.: ±10° Res.: 2° Range: 0–360° | Uncert.: ±15° Res.: 5° Range: 0–360° | As for wind speed regarding obstructions. Direction head to be set to True North not Magnetic North. |
|  | Accumulated precipitation | Uncert.: ±1 mm or 2% Res.: 0.1 mm Range: 0-1000 mm/day | Uncert.: ±3 mm or 5% Res.: 0.5 mm Range: 0-1000 mm/day | Uncert.: ±5 mm or 10% Res.: 1 mm Range: 0-500 mm/day | The area of the collector shall be at least 300 cm2. The typical gauge used in the region is 200 mm diameter (area 314 cm2). |
|  | Precipitation intensity | Uncert.: ±0.2 mm/h or 5% Res.: 0.1 mm/h Range: 0–500 mm/h | Uncert.: ±0.5 mm/h or 10% Res.: 0.3 mm/h Range: 0–300 mm/h | Uncert.: ±2 mm/h or 15% Res.: 1.0 mm/h Range: 0–300 mm/h | Must be suitable for heavy tropical downpours. Circular aperture on gauge should be minimum of 200 mm (AWS) |

### Atmosphere (Upper Air)

Table 17 Atmosphere (Upper Air) variable specification for each Tier (1-3). Tier 1 aligned with WMO No. 8 Annex 1.A and 1.G [1]. Tier 2 and 3 extrapolated from Tier 1 and informed by other sources, including expert judgement. Uncert. = Uncertainty; Res. = Resolution

| Category | Variable | Tier 1 | Tier 2 | Tier 3 | Comments |
| --- | --- | --- | --- | --- | --- |
| Atmosphere (Upper Air) | Atmospheric temperature (profile) | Uncert.: ±0.2°C Res.: Range: | Uncert.: ±0.5°C Res.: Range: | Uncert.: ±1.0°C Res.: Range: |  |
|  | Atmospheric pressure (profile) | Uncert.: ±0.2 hPa Range: | Uncert.: ±1.0 hPa Range: | Uncert.: ±2.0 hPa Range: | Balloon-borne or GNSS-derived |
|  | Relative humidity (profile) | Uncert.: ±2% Range: | Uncert.: ±5% Range: | Uncert.: ±10% Range: | Sensors must perform well in upper tropical troposphere |
|  | Wind/wind shear | Uncert.: ±0.5 m/s Direction: | Uncert.: ±1.0 m/s Direction: | Uncert.: ±3.5 m/s Direction: ± | From pilot balloon, radiosonde, or aircraft (AMDAR) |

### Marine

Table 18 Marine variable specification for each Tier (1-3). Class A aligned with WMO No. 8 Annex 1.A [1]. Tier 2 and 3 extrapolated from Tier 1 and informed by other sources, including expert judgement. Uncert. = Uncertainty; Res. = Resolution

| Category | Variable | Tier 1 | Tier 2 | Tier 3 | Comments |
| --- | --- | --- | --- | --- | --- |
| Marine | Sea surface temperature (SST) | Uncert.: ±0.2°C Res.: 0.1°C Range: -5 to +35°C | Uncert.: ±0.6°C Res.: 0.1°C Range: 0 to +35°C | Uncert.: ±1.0°C Res.: 0.1°C Range: 0 to +35°C | In situ thermistors on moorings or buoys |
|  | Coastal sea level (tide, surge, SSA) | Uncert.: ±0.05 m Res.: 0.01 m Sampling: ≤1 min | Uncert.: ±0.1 m Res.: 0.01 m Sampling: ≤1 min | Uncert.: ±0.5 m Res.: 0.01 m Sampling: ≤5 min | Radar or pressure tide gauges with cyclone resilience |
|  | Wave height (significant) | Uncert.: ±0.5 m for ≤5 m or 10% for >5 m Res.: 0.1 m Range: 0–50 m | Uncert.: ±1.0 m for ≤5 m or 15% for >5 m Res.: 0.1 m Range: 0–50 m | Uncert.: ±2.0 m for ≤5 m or 20% for >5 m Res.: 0.1 m Range: 0–50 m |  |
|  | Wave period (peak/mean) | Uncert.: ±0.5 s Res.: 1 s Range: 0-100 s | Uncert.: ±1 s Res.: 1 s Range: 0-100 s | Uncert.: ±2 s Res.: 1 s Range: 0-100 s |  |
|  | Wave direction (peak/mean) | Uncert.: ±10° Res.: 1° Range: 0–360° | Uncert.: ±15° Res.: 1° Range: 0–360° | Uncert.: ±20° Res.: 5° Range: 0–360° | Mooring-based wave buoys preferred |

### Freshwater & Hydrology

Table 19 Freshwater & Hydrology variable specification for each Tier (1-3). Tier 1 is aligned with National Environmental Monitoring Standards New Zealand (NEMS) QC600 standard [9-13]. Tier 2 and 3 are extrapolated from Tier 1 based on NEMS QC500 and QC400, as well as being informed by expert judgement. Uncert. = Uncertainty; Res. = Resolution, SpC = Specific Conductivity, DO = Dissolved Oxygen

| Category | Variable | Tier 1 | Tier 2 | Tier 3 | Comments |
| --- | --- | --- | --- | --- | --- |
| Freshwater & Hydrology | Water level (Rivers, Lakes) | Uncert.: The greater of ±3 mm or ±0.2% of effective water level. Res.: 1 mm Sampling interval: 5 min | Uncert.: The greater of ±3 mm or ±0.2% of effective water level. Res.: 3 – 5 mm Sampling interval: 5 min | Uncert.: The greater of ±3 mm or ±0.2% of effective water level. Res.: 3 – 5 mm Sampling interval: 5 min | Remote data retrieval; Secure mounting in flood-prone areas |
|  | Groundwater level | Uncert.: The greater of ± 10 mm or ± 0.2% of effective water level Res.: 1 mm Sampling interval: 15 min. | Uncert.: The greater of ±15 mm or ±0.5% of effective water level Res.: 3-5 mm Sampling interval: 15 -30 min. | Uncert.: The greater of ±20 mm or ±0.2% of effective water level Res.: 3-5 mm Sampling interval: 15-30 min | Borehole installations with vandal-proof casings |
|  | River discharge (Flow) | Uncert.: ±8% (discharge) | Uncert.: ±10% (discharge) | Uncert.: ±15% (discharge) | The in-situ measurement of water volume (flow) has a maximum permissible uncertainty of ±10% of reading over the entire range of expected flow rates to a confidence level of 95%. |
|  | Water quality (river and lake) | pH: ±0.2 pH units SpC: ±1 µS/cm or ±0.5% DO: ±0.3 mg/L | pH: ± pH units SpC: ± µS/cm or ± % DO: ± mg/L | pH: ± pH units SpC: ± µS/cm or ± % DO: ± mg/L | Remote logging preferred; Field calibration kits required |

### Land

Table 20 Land variable specification for each Tier (1-3). Tier 1 aligned with previous tender documents in the region, as well as National Environmental Monitoring Standards New Zealand [9-13]. Tier 2 and 3 extrapolated from Tier 1 and informed by expert judgement. Uncert. = Uncertainty; Res. = Resolution, VWC = Volumetric Water Content

| Category | Variable | Tier 1 | Tier 2 | Tier 3 | Comments |
| --- | --- | --- | --- | --- | --- |
| Land | Soil moisture | Uncert.: ±3% VWC Res.: ±1.5% | Uncert.: ±5% VWC Res.: ±1.5% | Uncert.: ±10% VWC Res.: ±1.5% | Can have multiple depths for sensors. Consider inclination, topography, soil type, depth of water table, obstacles and application needs (e.g., agricultural practices). |
|  | Ground (earth) temperature | Uncert.: ±0.3°C Res.: 0.1°C Range: 0 to +55°C | Uncert.: ±0.5°C Res.: 0.2°C Range: 0 to +55°C | Uncert.: ±1.0°C Res.: 0.5°C Range: 0 to +55°C | Standard soil depths (including some or all of 5, 10, 20, 50 and 100 cm). |
|  | Land use / land cover | Res.: 1-3 m (satellite/drone) | Res.: 10 m (satellite) | Res.: 30 m (satellite) | Ground truthing via GPS & drones. |

## Site & Installation Guidelines

Table 21 Site selection and installation guidelines for all tiers.

| Consideration | Specification |
| --- | --- |
| Exposure | Sensors positioned to avoid obstructions (wind, precipitation), representative of surrounding area. |
| Foundations | Secure, permanent mounting capable of withstanding cyclonic winds and flooding. |
| Cabling | UV-resistant, rodent-proof, with drip loops and surge protection. |
| Access | Safe, reliable access for maintenance and calibration. |
| Security | Fencing, signage, and where possible, remote monitoring to prevent vandalism/theft. |
| Environmental Impact | Minimal disturbance to local ecology; avoid culturally sensitive sites. |

# Procurement Tool Kit

Adopting a standardised procurement approach is essential to achieving regional interoperability, sustainability, and value for money. This section outlines a structured framework for procuring Earth observation systems in the Pacific, drawing on the principles of standardisation, whole-of-system thinking, and lifecycle management presented throughout this guidance.

A companion Procurement Toolkit has been developed alongside this Guidance Document, containing ready-to-use templates, evaluation matrices, and checklists. The toolkit is designed to:

- Ensure consistent and transparent procurement aligned with regional and WMO standards.

- Promote technical compatibility across the Pacific and with global systems (GBON, RBON, WIGOS, WIS 2.0).

- Support informed decision-making through a structured evaluation process.

- Enhance long-term sustainability by integrating whole-of-life considerations and “Supply & Support” contract models from the outset.

- Strengthen capacity development by selecting systems that support local maintenance and operations.

## Key Procurement Domains

The Procurement Tool Kit evaluates tenders across seven critical domains outlined in Table 22.

Table 22  Key procurement domains and considerations for evaluating tenders.

| Domain | Key Considerations |
| --- | --- |
| 1. System Design & Technical Suitability | Meets national use cases and regional standards; fit-for-purpose for Pacific conditions (cyclone-prone, high humidity, salt exposure, off-grid). |
| 2. Power, Telemetry & Data Integration | Solar resilience, battery backup, redundant communications, compatibility with national/regional data systems (e.g., CliDE, WIS 2.0). |
| 3. Installation & Localisation | Adherence to regional installation guidelines; involvement of local technical staff and NMHSs during installation. |
| 4. Vendor Capability & Support | Regional track record, after-sales support, availability of spares, warranties, training provisions, documentation quality. |
| 5. Lifecycle Costs & Maintenance | Total Cost of Ownership (TCO) evaluation covering purchase, installation, calibration, software, training, spare parts, and decommissioning. |
| 6. Sustainability & Environmental Protection | Materials durability, corrosion resistance, cyclone-proofing, end-of-life management, environmental safeguards. |
| 7. Data Sharing & Sovereignty | Supports national data ownership while capable of integration with regional/global platforms. Prefers open standards (BUFR, netCDF). Proprietary systems only if no open alternative exists, with guaranteed data export in open formats. |

## Evaluation Approach

The Procurement Toolkit uses scoring matrices to assess tenders objectively. Key features of the evaluation approach are outlined in Table 23.

Table 23 Example weighting scheme for procurement evaluation.

| Element | Description |
| --- | --- |
| Weighted Criteria | Domains are weighted based on national priorities, technical complexity, and regional integration needs (e.g., Vendor Support may carry higher weight than aesthetics). |
| Multidisciplinary Review | Evaluation panels should include NMHS technical staff, ICT specialists, procurement officers, and service delivery representatives. |
| Regional Alignment | Where possible, align evaluation criteria with neighbouring NMHSs to enable pooled procurement and cross-national compatibility. |
| Whole-of-Life Focus | Scoring emphasises long-term sustainability, not just upfront cost. “Supply & Support” models are incentivised. |

Example Weighting (Adjustable per NMHS):

- Technical Suitability: 25%

- Vendor Support & Track Record: 20%

- Lifecycle Cost & Sustainability: 20%

- Telemetry & Data Integration: 15%

- Installation & Training: 10%

- Environmental & Social: 10%

## Pre-Procurement Considerations

Before releasing tenders, NMHSs, national stakeholders, or regional agencies should take the necessary steps and actions to ensure a successful procurement outlined in Table 24.

Table 24 Pre-procurement steps and responsibilities.

| Step | Action |
| --- | --- |
| 1. Needs Validation | Confirm site-specific requirements, siting conditions, installation challenges, and power/telemetry constraints. |
| 2. Standards Alignment | Ensure alignment with WMO standards (WIGOS, GBON, RBON) and this Guidance Document’s tiered specifications. |
| 3. Stakeholder Engagement | Involve technical, procurement, and end-user representatives from relevant ministries/agencies. |
| 4. Regional Peer Review | Where appropriate, engage with SPREP, WMO, or other regional partners for tender content review. |
| 5. Budget Confirmation | Secure not only capital expenditure but also confirmed operational and maintenance funding for the system’s lifecycle. |
| 6. Capacity Readiness | Ensure staff training plans are in place and that national technical capacity matches system complexity. |

## “Supply & Support” Contracting Model

A strong recommendation is made to adopt “Supply & Support” contracts that extend beyond initial supply and installation to include long-term provisions. Minimum requirements are outlined in Table 25.

Key consideration: Minimum 3 to 5 years contract duration, align with expected calibration and major maintenance cycles.

Table 25 Minimum requirements for "Supply & Support" contracts models.

| Component | Minimum Requirements |
| --- | --- |
| Spare Parts | Guaranteed availability of critical spares for the system’s lifecycle, with regional stocking options. |
| Consumables | Supply of batteries, desiccants, filters, etc., for at least 3–5 years. |
| Technical Support | On-call remote support, with defined response times (e.g., ≤24 hours for critical failures). |
| Calibration Services | Access to calibration (onsite, regional, or return-to-factory) at agreed intervals. |
| Training | Comprehensive initial training plus periodic refresher courses, with materials in appropriate languages. |
| Software Updates | Updates, security patches, and compatibility maintenance for the contract duration. |

## Procurement Templates

The Procurement Toolkit includes pre-populated templates for major equipment types:

- Automatic Weather Station (AWS) – GBON/RBON compliant

- Climate Reference Station (Tier 1)

- Upper-Air Sounding System

- Marine Observation Buoy/Tide Gauge

- Hydrological Monitoring Station

- Telemetry & ICT Infrastructure Package

Each template includes:

- Technical specifications drawn from Section 5 of this guide

- Required documentation (manuals, calibration certificates, drawings)

- Training and support requirements

- Evaluation criteria and scoring guidance

## Post-Procurement Actions

Table 26 post-procurement actions and responsibilities.

| Phase | Key Activities |
| --- | --- |
| Contract Award | Notify all bidders, debrief unsuccessful vendors, sign contract with clear KPIs and reporting requirements. |
| Delivery & Inspection | Verify equipment against delivery notes, check for damage, validate calibration certificates. |
| Installation Oversight | NMHS staff should be present during installation; conduct joint commissioning tests. |
| Training & Handover | Ensure all planned training is delivered and documented; complete handover checklist. |
| Performance Monitoring | Track uptime, data quality, and support responsiveness against contract KPIs. |
| Contract Management | Schedule regular review meetings, maintain log of issues, plan for contract renewal/extensions. |

## Regional Pooled Procurement

Where feasible, NMHSs are encouraged to explore pooled procurement through regional mechanisms (e.g., SPREP, PIFS, WMO) to achieve:

- Economies of scale (bulk discounts)

- Standardisation across multiple countries

- Strengthened negotiating power with vendors

- Shared logistics and warehousing for spares

- Joint training and capacity building

SPREP/WRP can facilitate regional procurement coordination upon request.

# Spare Parts, Maintenance & Lifecycle Management

Effective observation networks require robust maintenance systems, reliable access to spare parts, and careful lifecycle management to ensure data quality and long-term sustainability. In the Pacific, where logistics, remoteness, and harsh environments create added challenges, proactive and strategic planning is not just beneficial but critical to the sustainability of these vital assets.

This section outlines a tiered approach to spare parts, maintenance, calibration, and lifecycle management, aligned with WMO guidelines and regional best practices. See Appendix A for a list of WMO Regional Instrument Centres (RICs) and Regional Training Centres (RTCs), as well as other regional partners, that can assist.

## Tiered Spare Parts Strategy

A tiered spare parts strategy is recommended to mitigate operational risks and minimise costly downtime (shown in Table 27).

Table 27 Tiered spare parts strategy, maintenance frequency, calibration verification frequency, and lifecycle expectations aligned with WMO Guidelines [31, 32].

| Tier | Key Spare Parts Strategy | Preventative Maintenance Frequency | Calibration Verification Frequency | Typical Lifecycle |
| --- | --- | --- | --- | --- |
| 1 | Full redundancy. Critical to have 100% sufficient functional spares/components on shelf. National inventory of all critical components. | Intensive: Monthly visual checks, quarterly detailed inspections with field verification, annual full service. | Annual (or per manufacturer, whichever is stricter) | 8-12 years with mid-life refurbishment |
| 2 | Essential spares for high-failure items. Regional pooling recommendations for expensive items. National stock of common spares. | Regular: Monthly to quarterly visual checks, bi-annual detailed inspections with field verification. | 1-2 years (aligned with regional RIC schedules) | 7-10 years with regular maintenance |
| 3 | Minimal spares. Focus on low-cost, replaceable sensors. Vendor-supported replacement. Regional pooling for common items. | As-needed / Reactive: Bi-annual checks, annual inspections with field verification. | 2-3 years or "as failed" | 5-8 years with lower-cost replacement strategy. |

## Spare Parts & Consumables Management

To ensure continuity of observations, each NMHS or relevant agency should maintain a three-level spare parts inventory outlined in Table 28.

Table 28 Three-level spare parts inventory structure (site, national, regional).

Key consideration: As a guiding principle, for critical Tier 1 (Reference) systems, a full replacement set of key components should be held in stock as a minimum requirement to safeguard the continuity of high-priority climate records and forecasting data.

| Level | Description | Examples |
| --- | --- | --- |
| Site-Level Stock | Minimum critical spares kit at each operational site for immediate repairs. | Fuses, cables, basic tools, desiccant, batteries. |
| National Central Warehouse | Consolidated national stock for higher-value, common-failure, and aging-risk components. | Sensors, data loggers, telemetry modules, solar panels. |
| Regional Procurement Agreements | Collaborative agreements with suppliers and partner agencies for rapid procurement of long-lead-time or catastrophic-fail items. | Specialised sensors, calibration equipment, cyclone-proof mounts. |

Through regional coordination:

- Stock essential spares for high-frequency replacement items (e.g., sensors, batteries, cabling, seals).

- Develop regional spare part pools to improve cost-effectiveness and reduce procurement delays, coordinated through WRP and SPREP.

- Ensure spare parts are standardised across systems to minimise stock requirements and simplify logistics.

## Maintenance Planning

Maintenance planning enables NMHSs to move from reactive repairs to scheduled, proactive maintenance. This is vital for data continuity, budgeting and resources. Maintenance planning requirements are outlined in Table 29.

Table 29 Maintenance planning requirements and components.

| Component | Requirement |
| --- | --- |
| National Maintenance Schedules | Develop and document schedules aligned with WMO standards and tier specifications [31, 32]. |
| Standard Operating Procedures (SOPs) | Include protocols for grounds maintenance (vegetation clearing), sensor cleaning, corrosion control, and cyclone preparedness. |
| Digital Maintenance Logs | Maintain asset history including calibration dates, parts replaced, faults, and performance trends. Use existing systems and tool to capture this information, e.g., CliDE metadata functionality. |
| Training | Train NMHS staff and local technicians in first-line maintenance, troubleshooting, and grounds maintenance, with support from RTCs and regional partners. |

## Field Verification & Calibration

### Field Verification

Field verification is the process of evaluating and confirming the accuracy, reliability, and operational performance of measuring instruments at their operational site against pre-defined requirements. This is carried out under actual environmental conditions using reference or portable standards.

Frequency:

- Tier 1: Quarterly

- Tier 2: Bi-annually

- Tier 3: Annually or as needed

### Calibration Infrastructure

Key consideration: The high cost and customs complexities of moving equipment across the Pacific for calibration should be addressed through intergovernmental arrangements to facilitate temporary import/export.

Access to calibration services is non-negotiable for data credibility. NMHSs should leverage RICs, coordinate with regional partners, and where capacity exists, develop in-country calibration capability for routine verifications.

## Traceability and Quality Assurance

The value of our Earth observation systems is realised only when the data they produce is reliable.

Traceability and Quality Assurance form the essential bridge between a sensor in the field and the confident decisions made by forecasters, climate scientists, and disaster managers. This process guarantees that our data is accurate, comparable with global networks, and fit for its purpose.

Guiding principles adopted in this document include:

- Ensure regular field verification of network instruments following recommend frequency in this guidance.

- Ensure calibration of field verification equipment using national, regional, or mobile calibration services following recommended frequency in this guidance.

- Access calibration services through RICs [34] (see Appendix A), which supports instrument calibration for the Pacific region, or through mobile calibration campaigns coordinated with regional partners.

Build in-country technician capacity with support from regional RTCs and/or partner-led training courses and workshops (see Appendix A).

Key consideration: WMO and the SOFF Secretariat have developed guidance on approaches for traceability assurance in GBON stations which is applicable to all network stations in the Region  ADDIN EN.CITE <EndNote><Cite><Author>SOFF</Author><Year>2025</Year><RecNum>43</RecNum><DisplayText>[33]</DisplayText><record><rec-number>43</rec-number><foreign-keys><key app="EN" db-id="5drtpt0wbe95whe50fb5rs0dvsrxsvdrz29t" timestamp="1764298210">43</key></foreign-keys><ref-type name="Generic">13</ref-type><contributors><authors><author>SOFF</author></authors></contributors><titles><title>SOFF Operational Guidance Handbook</title></titles><dates><year>2025</year></dates><urls><related-urls><url>https://www.un-soff.org/document/soff-operational-guidance-handbook/</url></related-urls></urls></record></Cite></EndNote>[33].

## Lifecycle Management

Procuring a sensor is a single event; managing the data it produces is a multi-decade responsibility. Lifecycle management is the practice of stewarding an asset from installation to decommissioning. It ensures the continuity of vital climate records, prevents catastrophic data gaps from unexpected failures, and safeguards our investments by planning for renewal before obsolescence strikes. Phases of the lifecycle management and recommended actions are outlined in Table 30.

Table 30 Lifecycle management phases and associated actions.

Key consideration: Technology evolves rapidly, especially in communications. Ensure compatibility with core regional systems like CliDE to avoid stranded data and duplicated infrastructure.

| Phase | Actions |
| --- | --- |
| Planning | Apply a Total Cost of Ownership (TCO) approach, accounting for procurement, operation, maintenance, calibration, and decommissioning. |
| Operation | Monitor performance, conduct regular maintenance, update software/firmware, and ensure cybersecurity. |
| Renewal | Plan for timely equipment upgrades to prevent obsolescence and ensure compatibility with evolving ICT/data standards. The importance of NMHS system compatibility is critical. |
| Decommissioning | Include a phase-out protocol with ≥12 months’ notice to ensure data overlap with new systems. Ensure proper disposal of hazardous materials. |

## Key Recommendations

- Adopt the tiered spare parts strategy outlined in Table 7.

- Develop and fund national maintenance plans with clear SOPs and digital logs.

- Leverage regional calibration networks and plan for logistics/customs facilitation.

- Implement lifecycle planning from procurement through to decommissioning.

- Build in-country technical capacity through training and mentorship programmes.

# ICT, Data Management, and Telemetry Considerations

Observation systems are only as valuable as the data they deliver. Pacific NMHSs must ensure reliable, secure, and wherever possible, interoperable ICT and telemetry systems that can transmit, process, and share data nationally, regionally, and globally. These systems should follow a “whole-of-system” approach from sensor to end-user.

## Telemetry Systems

### Redundant Communications Strategy

A single communication link is a single point of failure. The standard for critical sites (Tier 1 and 2) must be redundant communications. A primary/secondary model is the Pacific standard.

Table 31 Telemetry redundancy strategy options and technologies.

| Link Type | Technology Options | Use Case |
| --- | --- | --- |
| Primary | Cellular networks (GSM/3G/4G/5G) with industrial-grade modems and external antennas. | Where coverage exists; low cost, high bandwidth. |
| Secondary | Satellite services: Iridium SBD, Inmarsat BGAN, Starlink, VSAT. | Sites beyond cellular reach or as failover. Include pre-paid 2-year data plan in contract. |
| Tertiary (Optional) | Manual download via USB/SD card, or local wireless (Wi-Fi). | For non-critical sites or backup during extended outages. |

### Design for Resilience

Table 32 Design specifications for resilient telemetry and power systems.

| Consideration | Specification |
| --- | --- |
| Hardware Protection | Corrosion-resistant connectors and enclosures (IP68) for all outdoor telemetry hardware. |
| Data Storage | Data loggers must have sufficient on-board storage (minimum 1 year) to bridge extended telemetry outages. |
| Power Systems | Solar power systems must be sized to support both primary and secondary telemetry continuously, plus battery backup for ≥14 days. |
| Standardisation | Use the same make/model of loggers and telemetry units across the network where possible to simplify troubleshooting, spare parts, and training. |

## Data Management and Sovereignty

### National Data Sovereignty

The default should be that raw and primary processed data are stored on infrastructure owned or controlled by the NMHS, whether on national servers or a designated, secure cloud platform.

Table 33 Data storage options and sovereignty considerations.

Key consideration: All vendor contracts must explicitly state that the NMHS retains all rights to the data.

| Storage Option | Considerations |
| --- | --- |
| National Servers | Full control, higher upfront cost, requires IT maintenance and potentially requires high support from partners. |
| Designated Cloud Platform | Scalable, off-site redundancy, ongoing subscription costs. Must have data sovereignty clauses. |
| Hybrid Model | Critical data on national servers, archived/less critical data in cloud. |

### Robust Archiving and Metadata

Table 34 Archiving and metadata management requirements.

| Component | Requirement |
| --- | --- |
| Long-Term Archive | Separate from operational databases, with clear data preservation and backup policies. |
| Metadata | Collect and manage complete metadata (WIGOS Metadata Standard) at installation and update with every maintenance action. Essential for data discovery, use, and credibility. |
| Version Control | Maintain version histories of datasets and processing algorithms. |

## Integration and Interoperability

Data must flow seamlessly into the tools used by NMHSs for forecasting, warning, and service delivery.

### Direct Integration Pathways

Design systems so that data flows automatically from the telemetry gateway into key operational systems, such as:

Table 35 Primary data integration pathways into national, regional, and global systems.

| System Type | Examples |
| --- | --- |
| National Forecasting Platforms | Forecaster workstations, numerical weather prediction (NWP) models. |
| Data Management Systems | CliDE, WIS 2.0 nodes, national hydrological databases. |
| Early Warning Dissemination Systems | Common Alerting Protocol (CAP) platforms, SMS gateways, web portals. |

### Adhere to Already Adopted Standards

Table 36 Key standards for ensuring data interoperability.

| Standard | Purpose | Requirement |
| --- | --- | --- |
| WIS 2.0 | Global data exchange framework | All Tier 1 and 2 stations must be WIS 2.0 compliant. |
| BUFR/netCDF | Data formatting | Use for operational and archival data exchange. |
| Open APIs | System interoperability | Prefer systems with open, documented APIs for integration. |

## Cybersecurity and Data Integrity

### Secure Transmission and Access

Table 37 Cybersecurity measures for data transmission and access.

| Measure | Implementation |
| --- | --- |
| Encryption | Use HTTPS, VPNs, or encrypted satellite links for data transmission from field to server. |
| Authentication | Implement strong access controls (API keys, tokens, multi-factor authentication) for all system access. |
| Network Segmentation | Isolate operational networks from public-facing systems where possible. |

### Ensure Data Integrity and Availability

Table 38 Best practices for ensuring data integrity and availability.

| Practice | Description |
| --- | --- |
| Backup Regime | Daily incremental backups, weekly full backups, with regular tested recovery procedures. |
| Contingency Procedures | Fallback methods for critical data reception (e.g., manual satellite retrieval, redundant data paths). |
| Integrity Checks | Use checksums, hash validation, and automated data quality checks to detect corruption. |

## Data Sharing & Regional Collaboration

While NMHSs retain ownership, they should share data regionally and globally in line with the WMO Unified Data Policy, using standard formats and protocols.

Table 39 Mechanisms for regional and global data sharing.

Key consideration: Turns national observations into a global public good and increases their value through wider use in regional forecasting, research, and climate services.

| Mechanism | Purpose |
| --- | --- |
| WIS 2.0 | Primary channel for global and regional data exchange. |
| Regional Data Hubs | SPREP, SPC, or NMHS-led portals for Pacific-specific data sharing. |
| Bilateral Agreements | For tailored data exchanges with neighbouring countries or regional centres (RSMCs). |

## Capacity Development for ICT & Data Roles

Sustainable ICT and data management requires skilled personnel. Training should be integrated into the system lifecycle and tailored to specific roles (see also Section 9).

Table 40 ICT and data management roles and key competencies.

| Role | Key Competencies |
| --- | --- |
| ICT & Telemetry Engineer | Telecommunications protocols, solar power management, network security, hardware/software integration. |
| Data Management & QC Officer | Use of CliDE and other data systems, WIGOS metadata standards, data quality control procedures. |
| Forecasting & Warning Specialist | Application of real-time data in models, user-centric product development. |

## Key Recommendations

- Implement redundant telemetry for all Tier 1 and 2 stations.

- Ensure data sovereignty through contractual agreements and national/controlled storage.

- Adhere to WIS 2.0 and open standards for interoperability.

- Apply cybersecurity best practices including encryption, authentication, and regular backups.

- Share data regionally and globally to maximise value and support regional forecasting efforts.

- Build ICT and data management capacity through targeted training and regional partnerships.

# Capacity Development

A sustainable observation network is fundamentally dependent on a skilled, confident, and well-supported workforce. In the Pacific context, often characterised by small teams, remote locations, and high staff turnover, capacity development must be strategic, continuous, and institutionalised.

This section moves beyond one-off training events to provide a holistic framework for building human capability that aligns with the tiered observation system.

## Training Integrated with the System Lifecycle

Capacity development is not a standalone activity. It must be integrated into every stage of the observation system lifecycle, from procurement to decommissioning, and tailored to the specific requirements of each tier.

Table 41 Training framework aligned with the system lifecycle stages.

| System Lifecycle Stage | Capacity Development Focus | Example Activities |
| --- | --- | --- |
| Procurement & Design | Ensuring systems are supportable; defining training requirements in contracts. | Vendor questionnaires on training; mandating "Supply & Support" models with training clauses. |
| Installation & Commissioning | Hands-on, learning-by-doing for technical staff. | NMHS technicians working alongside vendor engineers; creating site-specific installation records. |
| Operations & Maintenance | Building routine and advanced competency for day-to-day sustainability. | Scheduled maintenance workshops; calibration training. |
| Data Management & Use | Maximising the value of observations for services. | Forecaster training on new data streams; data QC workshops; creating user-friendly products. |
| Refresh & Decommissioning | Managing technology change and preserving data heritage. | Training on new systems; data migration procedures; environmental disposal standards. |

## Role-Based Training Pathways & WMO Competency Alignment

Training must be targeted. The following framework defines core roles within the observation network and maps them to the required competencies and primary tier focus.

Table 42 Role-based training pathways aligned with WMO competencies and tiered systems.

| Role | Primary Tier Focus | Core Responsibilities | Key Training & Competency Requirements |
| --- | --- | --- | --- |
| Observation Technician | All Tiers | Routine maintenance, visual checks, basic sensor replacement, groundskeeping, simple troubleshooting. | Vendor-specific equipment training; Safety procedures for remote / coastal sites |
| Observation Systems Specialist | Tier 1, Tier 2 | Advanced diagnostics, calibration, network performance analysis, managing technician teams, complex repairs. | Meteorology and traceability principles; Data quality control procedures |
| Data Management & QC Officer | All Tiers | Data ingestion, archival, quality control, metadata management, application of QC flags, user support. | Use of CliDE and other data systems; Understanding of WIGOS metadata standards |
| Forecasting & Warning Specialist | Tier 2 | Interpreting real-time observations for forecasts and warnings; providing feedback on data quality and gaps. | Application of tiered data in models; User-centric product development |
| ICT & Telemetry Engineer | All Tiers | Managing telemetry networks, data flow architecture, cybersecurity, hardware/software integration. | Telecommunications protocols (WIS 2.0); Solar power system management |
| Observation Network Manager | All Tiers | Strategic planning, lifecycle budgeting, vendor management, overall network performance and reporting. | Procurement and contract management; Performance monitoring and evaluation |

## Leveraging and Strengthening the Regional Training Ecosystem

The Pacific benefits from a network of Regional Training Centres (RTCs) and technical partners. NMHSs should actively plan to utilise and contribute to this ecosystem.

- Align national training needs with other regional partners.

- Invest in developing in-country and regional trainers to ensure sustainability and local relevance of training.

- Participate in and contribute to regional forums, workshops, and online communities of practice for technicians, data managers, and forecasters.

## Mentorship & Succession Planning

To mitigate the impact of staff turnover and preserve critical institutional knowledge:

Table 43 Mentorship and succession planning strategies.

| Strategy | Implementation |
| --- | --- |
| Structured Mentorship | Pair junior staff with experienced mentors within the NMHS or from a regional partner agency. |
| Succession Planning | For roles essential to Tier 1 and critical Tier 2 operations, identify and actively develop at least one successor through planned experience and training. |
| Knowledge Repositories | Systematically document SOPs, site-specific knowledge, troubleshooting guides, and configuration details in a central, accessible digital repository. |

## Mandating Training in Procurement

Training must be a non-negotiable component of all equipment and system procurement.  All procurement contracts, especially under "Supply & Support" models, must include:

Table 44 Mandatory training and documentation requirements in procurement contracts.

Key consideration: A vendor’s proven ability and commitment to training Pacific Island staff should be a heavily weighted criterion in procurement evaluations.

| Requirement | Details |
| --- | --- |
| Initial Training | Comprehensive hands-on training during installation / commissioning for all relevant staff. |
| Documentation | Clear, durable documentation (hard and soft copy) in appropriate languages. |
| Refresher Options | Availability of refresher or advanced training sessions during the contract period. |
| Remote Support | Access to remote technical assistance and troubleshooting guidance. |

## Monitoring the Impact of Capacity Development

Investment in people must be measured by its results. NMHSs should track:

Table 45  Metrics for monitoring the impact of capacity development.

| Metric | Method |
| --- | --- |
| Training Participation | Number of staff trained, days of training, certifications achieved. |
| Performance Improvement | Station uptime, data quality scores, meantime-to-repair, reduction in repeat faults. |
| Service Impact | User feedback on forecasts/warnings; evidence of improved data use in decision-making. |
| Institutional Strength | Staff retention rates, progression into senior roles, contribution to regional training. |

## Cross-Training & Regional Collaboration

Table 46 Cross-training and regional collaboration initiatives.

| Initiative | Description |
| --- | --- |
| Cross-Training | Train observation staff in basic telemetry troubleshooting and ICT staff in meteorological data fundamentals to build a more resilient, multi-skilled team. |
| Regional Partnerships | Actively participate in and contribute to regional training offered by RTCs and partner organisations. |
| Knowledge Sharing | Encourage sharing of tools, scripts, configuration templates, and SOPs between Pacific NMHSs to avoid “reinventing the wheel.” |
| Train-the-Trainer Programmes | Develop in-country trainers to ensure sustainability and local relevance of training. |

## Funding & Resourcing Capacity Development

Capacity development must be budgeted for explicitly within both project and operational funding.

Table 47 Funding and resourcing sources for capacity development.

| Source | Potential Use |
| --- | --- |
| Project Budgets | Include training and capacity-building line items in all observation system procurement and upgrade projects. |
| National Budgets | Advocate for dedicated training funds within NMHS operational budgets. |
| Donor Funding | Seek aligned donor support for regional training initiatives, scholarships, and train-the-trainer programmes. |
| In-Kind Support | Leverage partner agency support for mentorship, remote training, and materials development. |

## Key Recommendations

- Adopt role-based training pathways aligned with WMO competencies and tiered system requirements.

- Integrate training into all procurement contracts, mandating vendor-provided training and documentation.

- Leverage the regional training ecosystem through RTCs, USP, SPC, and technical partners.

- Implement mentorship and succession planning to retain institutional knowledge.

- Monitor and evaluate training impact on both technical performance and service delivery.

- Budget for capacity development as a core component of sustainable observation networks.

# Governance, Implementation & Regional Coordination

The successful adoption and sustainability of this guidance require clear roles and accountability across regional and national levels.

## Roles and Responsibilities

Table 48 Roles and responsibilities of key actors in the implementation of this guidance.

Key consideration: WRP support for NMHS operations will be prioritised for those adopting and complying with this guidance. Non-compliance may limit access to WRP technical and financial assistance.

| Actor | Primary Role | Key Responsibilities |
| --- | --- | --- |
| WRP Programme (through SPREP) | Regional oversight & adoption | Endorse and promote the guidance; coordinate regional compliance; facilitate pooled procurement; provide technical review of investments. |
| Pacific NMHSs | National implementation | Adopt standards in national planning; apply tiered specifications in procurement; ensure ongoing maintenance and data quality; report on performance. |
| Development Partners & Donors | Alignment & funding | Ensure investments align with this guidance; support WRP-coordinated procurement; avoid fragmented or incompatible systems. |
| Regional Technical Partners | Technical support & capacity building | Provide calibration, training, and maintenance support in line with tiered specifications; participate in regional review panels. |

## Compliance Monitoring & Assurance

To ensure consistent application of standards, the following mechanisms will be used:

- WRP staff and SPREP officers will review major procurement plans (>USD 20,000) to ensure alignment with guidance before tenders are issued.

- Where requested, SPREP can conduct pooled or delegated procurement on behalf of NMHSs, ensuring full compliance with regional standards.

- NMHSs will complete a brief annual report against guidance KPIs (e.g., % of stations by tier, data availability, maintenance logs).

- Every 2 years, WRP/SPREP will coordinate light-touch audits of a sample of NMHSs to review compliance and identify support needs.

## Phased Implementation Timeline

Implementation should be phased to allow for capacity building, system alignment, and sustainable investment:

Table 49 Phased implementation timeline for adopting the guidance.

| Phase | Timeline | Key Activities |
| --- | --- | --- |
| Phase 1: Awareness & Adoption | Years 1–2 | Launch guidance; conduct regional workshops; integrate into WRP partner agreements; NMHSs map existing stations to tiers. |
| Phase 2: Procurement Alignment | Years 2–4 | Apply tiered specs to all new procurements; establish regional review panels; begin pooled procurement pilots; update national maintenance plans. |
| Phase 3: Full Compliance & Sustainability | Years 4–6 | All new investments fully compliant; regional spare parts pool operational; NMHSs report annually on performance; guidance updated based on lessons learned. |
| Phase 4: Regional Integration & Scaling | Years 6–10 | Full interoperability achieved; data sharing seamless; regional calibration and training hubs fully functional; guidance embedded in Pacific meteorological strategy. |

## Governance Structure for Regional Coordination

A practical governance structure is recommended to oversee implementation (see Figure 3).

Figure 3 Governance and coordination structure for the implementation of this guidance across the Pacific region, showing roles of WRP, NMHSs, development partners, and regional technical bodies.

# Acknowledgements

The development of this Guidance Document has been a collaborative effort, made possible through the contributions, expertise, and support of numerous individuals and organisations across the Pacific region and beyond. The authors extend their sincere gratitude to all who participated in its creation.

###### Weather Ready Pacific (WRP) Programme

We thank the WRP Programme for the vision, funding, and regional mandate to develop this guidance, and for their commitment to strengthening Pacific resilience through improved Earth observation systems.

###### Pacific National Meteorological and Hydrological Services (NMHSs)

We acknowledge the directors, technical staff, and managers of Pacific NMHSs who provided invaluable insights, real-world challenges, and operational perspectives through workshops, reviews, and bilateral consultations. Your practical experience is the foundation of this document.

###### Regional Organisational Partners

Secretariat of the Pacific Regional Environment Programme (SPREP) – For regional coordination, hosting consultations, and ongoing stewardship of the WRP Programme.

World Meteorological Organization (WMO) – For technical guidance, alignment with global standards (WIGOS, GBON, RBON), and facilitation of expert review.

Pacific Community (SPC) – For contributions in hydrology, GIS, ICT, and capacity development.

###### Technical & Development Partners

Bureau of Meteorology (BoM), Australia – For technical input, calibration support, and training expertise.

Japan Meteorological Agency (JMA) – For contributions to standards, training, and regional capacity support.

Earth Sciences New Zealand & MetService New Zealand – For hydrological and meteorological technical guidance, training, and Pacific project experience.

Systematic Observations Financing Facility (SOFF) – For alignment with GBON implementation and sustainability frameworks.

Intra-ACP Climate Services and Related Applications (CLIMSA) Programme – For foundational work on regional observation standards.

###### Workshop Participants & Reviewers

We thank the participants of the Christchurch Workshop (November 2025) and other regional consultations for their constructive feedback, which significantly shaped the tiered approach and practical recommendations contained in this guide.

###### Disclaimer

This document has been prepared for the Weather Ready Pacific Programme. The views and recommendations expressed are those of the authors and do not necessarily reflect the official policy or position of all participating organisations. While every effort has been made to ensure accuracy, the authors and supporting organisations accept no liability for any errors, omissions, or consequences arising from the use of this guidance.

###### How to Cite This Document

O’Neill, S., Elley, G., Harper, A., & Powers-Tora, M. (2026). Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands. Prepared for the Weather Ready Pacific (WRP) Programme. Wellington, New Zealand.

# Glossary of terms and acronyms

## Terms and Definitions

| Term | Definition | Source(s) |
| --- | --- | --- |
| Air Temperature | The temperature of the air, typically measured in a shelter 1.25–2 m above ground level. Maximum, minimum, and mean temperatures are the highest, lowest, and average values over a specified period (e.g., 24 hours). | [1] |
| Relative Humidity | The ratio, in percent, of the observed vapour pressure to the saturation vapour pressure with respect to water at the same temperature and pressure. | [1] |
| Air Pressure | The force exerted by the atmosphere per unit area. Also called atmospheric pressure. Expressed in hectopascals (hPa). | [1] |
| Wind Speed and Direction | Wind Speed: The rate of motion of air. Wind Direction: The direction from which the wind blows, expressed in degrees from true north. | [1] |
| Wind Gust | The maximum value of the wind speed averaged over a short time interval, typically 3 seconds. | [1] |
| Accumulated Precipitation (total amount of precipitation) | The total amount of precipitation collected over a specified period. Expressed as the depth of water in millimetres (mm). | [1] |
| Precipitation Intensity | The rate of fall of precipitation, expressed in millimetres per hour (mm/h). | [1] |
| Atmospheric Temperature (Profile) | The temperature of the air at various levels in the atmosphere, measured by radiosondes, aircraft, or other profiling systems. | [1] |
| Atmospheric Pressure (Profile) | The atmospheric pressure at various levels above the surface. | [1] |
| Relative Humidity (Profile) | The relative humidity at various levels in the atmosphere. | [1] |
| Wind/Wind Shear (Profile) | Wind: Speed and direction at altitude. Wind Shear: The variation in wind speed and/or direction over a short distance, primarily vertical. | [1]; [35] |
| Dew Point Temperature | Temperature to which a volume of air must be cooled at constant pressure and constant moisture in order to reach saturation; any further cooling causes condensation. | [35] |
| Sea Surface Temperature (SST) | The temperature of the water layer from 1 mm to 20 m below the sea-air interface. An Essential Climate Variable (ECV). | [36] |
| Coastal Sea Level | The height of the sea surface relative to a fixed benchmark on land. Includes components from tides (astronomical), storm surge (meteorological), and long-term anomalies. | [36] |
| Sea State | A description of the ocean surface with respect to wind waves and swell. Characterised by significant wave height (Hs), wave period, and wave direction. | [35] |
| River Height / Discharge | River Height (Stage): The elevation of the water surface at a gauging station. Discharge: The volume of water flowing through a cross-section per unit of time (m³/s). | [37]; [35]; [36] |
| Groundwater Levels | The depth of the water table, or the hydraulic head, in an aquifer, measured relative to a datum (often mean sea level). | [35]; [36] |
| Water Quality | The physical, chemical, and biological characteristics of water, determining its fitness for various uses and ecosystem health. | [38] |
| Soil Moisture | The average water content in the soil, which can be expressed in volumetric, gravimetric, or relative units. An Essential Climate Variable (ECV). Expressed as Volumetric Water Content (VWC). | [36] |
| Soil Temperature | The temperature of the soil at a specified depth. The standard depths for soil temperature measurements are 5, 10, 20, 50, and 100 cm below the surface. | [1]; [36] |
| Land Use / Land Cover | Land Cover: The observed (bio)-physical cover of the earth's surface for regional and global climate applications. Land Use: The arrangements, activities, and inputs people undertake in a certain land cover type. | [36] |
| Water Level | The elevation of the water surface of a river (or a lake, reservoir) regarding a reference (the ellipsoid). | [36] |
| Effective Water Level | (Operational term) The difference between the measured water level and the anticipated minimum water level over the station's lifecycle. Used for defining required uncertainty. | Derived from [9]; [36] |
| Interoperability | The ability of diverse systems and organisations to work together (inter-operate). In meteorology, this enables seamless data exchange. | [6] |
| Total Cost of Ownership (TCO) | A financial estimate to determine the direct and indirect costs of a product or system over its entire lifecycle. | [16] |
| Observation Network | A coordinated set of observing stations and facilities designed to monitor the state of the environment. | Derived from [1] |
| Telemetry | The automated process of measuring and transmitting data from remote sources. | Derived from [1] |
| Calibration | Operation that, under specified conditions, in a first step, establishes a relation between the quantity values with measurement uncertainties provided by measurement standards and corresponding indications with associated measurement uncertainties and, in a second step, uses this information to establish a relation for obtaining a measurement result from an indication. | [1] |
| Maintenance Planning | The scheduled and systematic activities required to ensure an observing system continues to perform within its specified requirements. | Derived from [1] |

## Acronyms and Abbreviations

| Acronym | Full Form | Notes / Context |
| --- | --- | --- |
| ACP | African, Caribbean and Pacific | Group of countries engaged in the EU-funded Intra-ACP programmes (e.g., CLIMSA). |
| AWS | Automatic Weather Station | Unattended meteorological station measuring key weather variables automatically. |
| BoM | Bureau of Meteorology (Australia) | Key regional partner supporting Pacific NMHSs. |
| BUFR | Binary Universal Form for the Representation of meteorological data | WMO standard for transmitting observational data in binary format. |
| CLIMSA | Intra-ACP Climate Services and Related Applications Programme | EU-funded programme strengthening climate services in ACP regions. |
| COSPPac | Climate and Oceans Support Program in the Pacific | Australian and New Zealand Government-funded programme enhancing Pacific climate services. |
| CSV | Comma-Separated Values | Simple text file format used for data exchange and storage. |
| DIAL | Differential Absorption Lidar | Remote sensing instrument for atmospheric gas concentration measurements. |
| DO | Dissolved Oxygen | Key variable for freshwater and marine environmental monitoring. |
| DRR | Disaster Risk Reduction | Framework for reducing exposure and vulnerability to natural hazards. |
| EW4ALL | Early Warnings for All | Global UN initiative to ensure all people are protected by early warning systems by 2027. |
| GBON | Global Basic Observing Network | WMO initiative ensuring sustained, high-quality global meteorological observations. |
| GNSS | Global Navigation Satellite System | Used for atmospheric profiling, timing, and positioning of observation stations. |
| GPRS | General Packet Radio Service | Mobile data communication used by some remote telemetry systems. |
| HF | High Frequency | Radio band used for long-distance communication from remote observation sites. |
| ICT | Information and Communication Technology | Encompasses digital systems for data management, processing, and sharing. |
| JSON | JavaScript Object Notation | Lightweight data format used in web applications and APIs. |
| M&E | Monitoring and Evaluation | Used for tracking performance and impact of projects and systems. |
| NEMS | National Environmental Monitoring Standards (New Zealand) | Reference standards for environmental monitoring methods and QA/QC. |
| NMHS | National Meteorological and Hydrological Service | National authority responsible for weather, climate, and hydrology observations. |
| O&M | Operations and Maintenance | Activities ensuring continuous and reliable performance of observation systems. |
| PAGASA | Philippine Atmospheric, Geophysical, and Astronomical Services Administration | Able to provide regional support and training through RIC and RTC. |
| PICI | Pacific Island Communications, Infrastructure Panel | Regional mechanism supporting communications standards and infrastructure. |
| PICTs | Pacific Island Countries and Territories | Refers collectively to the nations and territories of the Pacific region. |
| PIFS | Pacific Islands Forum Secretariat | Regional intergovernmental organisation coordinating political and policy cooperation. |
| QC | Quality Control | Procedures ensuring Uncertainty and reliability of observational data. |
| RA V | WMO Regional Association V (South-West Pacific) | WMO regional grouping for the Pacific and parts of Southeast Asia. |
| RBON | Regional Basic Observing Network | WMO regional network supporting GBON implementation and regional data exchange. |
| RIC | Regional Instrument Centre | WMO-designated facility supporting calibration, maintenance, and training. |
| RO | Radio Occultation | Satellite-based technique for atmospheric profiling using GNSS signals. |
| RSMC | Regional Specialized Meteorological Centre | WMO-designated centres providing regional forecasting and analysis (e.g., RSMC Nadi). |
| RTC | Regional Training Centre | WMO-accredited institution providing technical training for NMHSs. |
| SOFF | Systematic Observations Financing Facility | Global financing mechanism supporting GBON implementation in developing countries. |
| SPC | Pacific Community | Regional scientific and technical organisation supporting development across PICTs. |
| SPREP | Secretariat of the Pacific Regional Environment Programme | Regional intergovernmental organisation focused on environmental protection and climate resilience. |
| SSA | Sea Surface Anomaly | Deviation of sea surface variables (e.g., temperature or height) from the mean. |
| SST | Sea Surface Temperature | Key climate indicator monitored regionally and globally. |
| TCO | Total Cost of Ownership | Full lifecycle cost including procurement, maintenance, training, and disposal. |

# References

1.	WMO, Guide to Instruments and Methods of Observation (WMO-No. 8). 2025. Available from: https://library.wmo.int/idurl/4/41650

2.	WMO. Early Warnings for All (EW4All) 2025  [cited 2025 30 October]; Available from: https://earlywarningsforall.org.

3.	WMO. Systematic Observations Financing Facility (SOFF). 2025  [cited 2025 30 October]; Available from: https://wmo.int/activities/systematic-observations-financing-facility-soff.

4.	ClimSA. Intra-ACP Climate Services and Related Applications (CLIMSA) Programme. 2025  [cited 2025 30 October]; Available from: https://www.climsa.org.

5.	BoM. Climate and Oceans Support Program in the Pacific (COSPPac). 2025  [cited 2025 30 October]; Available from: https://www.bom.gov.au/government-and-industry/government-sectors/indo-pacific-development/climate-and-oceans-support-program-in-the-pacific.

6.	WMO, WMO Unified Data Policy. 2022. Available from: https://library.wmo.int/idurl/4/58009.

7.	WMO, Manual on the WMO Integrated Global Observing System (WMO-No. 1160). 2024. Available from: https://library.wmo.int/viewer/55063/?offset=#page=1&viewer=picture&o=bookmark&n=0&q=.

8.	WMO, Guide to the WMO Integrated Global Observing System (WMO-No. 1165). 2024. Available from: https://library.wmo.int/viewer/55696/?offset=#page=1&viewer=picture&o=bookmark&n=0&q=.

9.	NEMS, Open Channel Flow Measurement: Measurement, Processing and Archiving of Open Channel Flow Data. 2013, National Environmental Monitoring Standards. Available from: https://bucketeer-54c224c2-e505-4a32-a387-75720cbeb257.s3.amazonaws.com/public/Documents/NEMS-Open-Channel-Flow-Measurement-v1.1.pdf.

10.	NEMS, Water Quality, Part 1 of 4: Sampling, Measuring, Processing and Archiving of Discrete Groundwater Quality Data. 2019, National Environmental Monitoring Standards. Available from: https://bucketeer-54c224c2-e505-4a32-a387-75720cbeb257.s3.amazonaws.com/public/Documents/NEMS-Water-Quality-Part-2-Sampling-Measuring-Processing-and-Archiving-of-Discrete-River-Water-Quality-Data-v1.0.0.pdf.

11.	NEMS, Water Quality Part 2 of 4: Sampling, Measuring, Processing and Archiving of Discrete River Water Quality Data. 2019, National Environmental Monitoring Standards. Available from: https://bucketeer-54c224c2-e505-4a32-a387-75720cbeb257.s3.amazonaws.com/public/Documents/NEMS-Water-Quality-Part-2-Sampling-Measuring-Processing-and-Archiving-of-Discrete-River-Water-Quality-Data-v1.0.0.pdf.

12.	NEMS, Water Quality Part 3 of 4: Sampling, Measuring, Processing and Archiving of Discrete Lake Water Quality Data. 2019, National Environmental Monitoring Standards. Available from: https://bucketeer-54c224c2-e505-4a32-a387-75720cbeb257.s3.amazonaws.com/public/Documents/NEMS-Water-Quality-Part-3-Sampling-Measuring-Processing-and-Archiving-of-Discrete-Lake-Water-Quality-Data-v1.0.0.pdf.

13.	NEMS, Continuous Measurement of Water Levels for Environmental Monitoring Purposes. 2025, National Environmental Monitoring Standards. Available from: https://bucketeer-54c224c2-e505-4a32-a387-75720cbeb257.s3.amazonaws.com/public/Documents/Water-Level-v4.0.1.pdf.

14.	WMO, Guidelines on the Role, Operation and Management of National Meteorological and Hydrological Services (WMO-No. 1195). 2017. Available from: https://library.wmo.int/viewer/55823/?offset=#page=1&viewer=picture&o=bookmark&n=0&q=.

15.	WMO, Compendium of WMO Competency Frameworks (WMO-No. 1209). 2019. Available from: https://library.wmo.int/viewer/56877/?offset=#page=1&viewer=picture&o=bookmark&n=0&q=.

16.	Bank, W., Charting a Course for Sustainable Hydrological and Meteorological Networks in Developing Countries. 2022. Available from: https://www.gfdrr.org/en/publication/charting-course-sustainable-hydromet-observation-networks.

17.	WMO, WMO Region Association V (RA-V) RBON Workshop - DRAFT. 2025. Available from: https://wmoomm.sharepoint.com/sites/wmocpdb/eve_group/Forms/AllItems.aspx?id=%2Fsites%2Fwmocpdb%2Feve%5Fgroup%2FExpert%20Team%20on%20Earth%20Observing%20System%20Design%20and%20Evolution%5F1c36026f%2Dba58%2Def11%2Dbfe3%2D000d3aa8d380%2FGroup%20Members%2FWorkplan%2FRBON%2FRA%2DV%2FRA%2DV%2DRBON%2Dworkshop%2FWorkshop%20report%2FRAV%5FRBON%5FWorkshop%5F2025%5Ffinal%5Freport%2Epdf&parent=%2Fsites%2Fwmocpdb%2Feve%5Fgroup%2FExpert%20Team%20on%20Earth%20Observing%20System%20Design%20and%20Evolution%5F1c36026f%2Dba58%2Def11%2Dbfe3%2D000d3aa8d380%2FGroup%20Members%2FWorkplan%2FRBON%2FRA%2DV%2FRA%2DV%2DRBON%2Dworkshop%2FWorkshop%20report&p=true&ga=1.

18.	PMC. The Pacific Island Communications, Infrastructure (PICI) Panel. 2025  [cited 2025 30 October]; Available from: https://www.pacificmet.net/pmc/expert-panels/pici.

19.	WMO, Instruments and Observing Methods - Guidance on Automatic Weather Systems and their Implementation (WMO/TD-No. 862), O.E.B.A.B.o. Meteorology, Editor. 1997. Available from: https://library.wmo.int/viewer/41844/?offset=#page=1&viewer=thumb&o=bookmarks&n=0&q=.

20.	Intech. Pt100 Patch – Surface Pt100 Patch. 2025  [cited 2025 30 October]; Available from: https://www.intech.co.nz/product/pt100-patch/.

21.	OTT. OTT Pluvio² L – Weighing Rain Gauge. 2025  [cited 2025 30 October]; Available from: https://www.ott.com/products/meteorological-sensors-26/ott-pluvio2-l-weighing-rain-gauge-1521/productAction/outputAsPdf/.

22.	OTT. SUTRON Tipping Bucket Rain Gauge with Stainless Steel Tipping Mechanism. 2025  [cited 2025 30 October]; Available from: https://www.otthydromet.com/en/p-sutron-rain-gauge-stainless-steel-tipping-bucket-001-inchtip50-ft15-m-cable-with-stainless-steel-tipping-mechanism/5600-0425-6

23.	OTT. Orpheus Mini. 2025  [cited 2025 30 October]; Available from: https://www.ott.com/download/leaflet-groundwater-datalogger-ott-orpheus-mini-1/.

24.	Scientific, C. CS106 Barometer. 2025  [cited 2025 30 October]; Available from: https://www.campbellsci.com/cs106.

25.	Scientific, C. CS655 12 cm Soil Moisture and Temperature Sensor. 2025  [cited 2025 30 October]; Available from: https://www.campbellsci.com/cs655.

26.	Scientific, C. 109SS Stainless-Steel Temperature Probe for Harsh Environments. 2025  [cited 2025 30 October]; Available from: https://www.campbellsci.com/109ss.

27.	Setra. Setra Model 278 – Product Data Sheet and Application Notes. 2025  [cited 2025 30 October]; Available from: https://www.setra.com/hubfs/Product_Data_Sheets/Setra_Model_278_Data_Sheet.pdf.

28.	Solutions, R.M. High-End Meteorological Probe - Rotronic HC2A-S3A High Accuracy Humidity and Temperature Measurement. 2025  [cited 2025 30 October]; Available from: https://www.processsensing.com/en-us/products/humiditiy-probes-hc2a-s3a.htm.

29.	Unidata. Linear, Thermistor and SDI-12 Temperature Probes. 2025  [cited 2025 30 October]; Available from: https://www.unidata.com.au/products/environmental-monitoring-modules/thermistor-temperature-probes/.

30.	Vaisala. Product Data Sheets and Application Notes. 2025  [cited 2025 30 October]; Available from: https://www.vaisala.com/en/search?k=&content_group=products+and+services.

31.	WMO, Challenges in the Transition from Conventional to Automatic Meteorological Observing Networks for Long-term Climate Records (WMO-No. 1202). 2017. Available from: https://library.wmo.int/viewer/55812/?offset=#page=1&viewer=picture&o=bookmark&n=0&q=.

32.	WMO, Guidelines on Surface Station Data Quality Control and Quality Assurance for Climate Applications (WMO-No. 1269). 2021. Available from: https://library.wmo.int/viewer/57727/?offset=#page=1&viewer=picture&o=bookmark&n=0&q=.

33.	SOFF, SOFF Operational Guidance Handbook. 2025. Available from: https://www.un-soff.org/document/soff-operational-guidance-handbook/.

34.	WMO. Regional Instrument Centres. 2025  [cited 2025 28th November]; Available from: https://wmo.int/regional-instrument-centres.

35.	WMO, International meteorological vocabulary. 1992, WMO: Geneva. p. 802. Available from: https://library.wmo.int/idurl/4/35809.

36.	C3S, W.I.-U.I.U., The 2022 GCOS ECVs Requirements 2022 edition - Updated in 2025, in GCOS-245. 2025, WMO: Geneva. Available from: https://library.wmo.int/idurl/4/58111.

37.	WMO, Guide to Hydrological Practices. 2009, WMO: Geneva. Available from: https://library.wmo.int/idurl/4/36066.

38.	WHO, Guidelines for drinking-water quality: small water supplies. 2024: Geneva. Available from: https://iris.who.int/server/api/core/bitstreams/4a3dc8a4-9f08-4d62-b7ec-96d70be6b177/content.

- Regional Support for Calibration, Training, and Capacity Development in the Pacific

Part A: WMO Regional Instrument Centres (RICs) and Regional Training Centres (RTCs)

| Centre | Location | Type | Core Functions | Relevance for Pacific NMHSs |
| --- | --- | --- | --- | --- |
| RIC Manila | Philippines (Philippine Atmospheric, Geophysical, and Astronomical Services Administration) | RIC | To assist WMO Members in calibrating their NMHS measurement reference standards (or field verification equipment in case a NMHS does not have measurement reference standards) and related monitoring instruments. To make a positive contribution regarding quality of observations. Advise Members on enquiries regarding instrument performance, maintenance and the availability of relevant guidance materials. To lead, actively participate in, or assist, in the organization of workshops and training events on calibration and maintenance of instruments. | Able to provide calibration and regional support for Pacific NMHS. Able to calibrate: - Temperature - Relative humidity - Atmospheric pressure - Wind - Precipitation |
| RIC Melbourne | Australia (Bureau of Meteorology) | RIC | As above. | Able to provide calibration and regional support for Pacific NMHSs until RIC Nadi is operational Able to calibrate: - Temperature - Relative humidity - Atmospheric pressure - Precipitation |
| RIC Tsukuba | Japan (Japan Meteorological Agency) | RIC | As above. | Able to provide calibration and regional support for Pacific NMHSs; potential partner for specialised calibration needs. Able to calibrate: - Temperature - Relative humidity - Atmospheric pressure - Wind |
| RIC Nadi (planned 2027) | Fiji (SPREP/WMO partnership) | RIC | As above. | Proposed to calibrate: - Temperature - Relative humidity - Atmospheric pressure |
| RTC Nadi (planned 2027) | Fiji (in collaboration with USP/SPREP) | RTC | Training of Pacific NMHS staff in instrument maintenance, data management, ICT, and forecasting | Will be the dedicated Pacific training hub, building sustainable regional workforce capacity |
| RTC Melbourne | Australia (Bureau of Meteorology) | RTC | Training on forecasting, climate services, instrument management, and ICT systems | Current main training provider for Pacific NMHSs; will complement RTC Nadi from 2027 |
| RTC Tsukuba | Japan (Japan Meteorological Agency) | RTC | Advanced training in meteorology, instrument use, and observation techniques | Specialised training opportunities for Pacific NMHS staff |

Part B: Regional Partners Supporting Training, Calibration, and Service Delivery

| Organisation | Location | Role | Core Functions |
| --- | --- | --- | --- |
| SPREP (Secretariat of the Pacific Regional Environment Programme) | Apia, Samoa | Regional Organisation | Training, technical support for climate services |
| USP (University of the South Pacific) | Suva, Fiji | Academic Partner | Provides formal education and vocational training in meteorology, ICT, and climate science; role in RTC Nadi |
| Earth Sciences New Zealand | Wellington & Christchurch, New Zealand | Technical Partner | Hydrology, marine, and climate services; training, maintenance, and calibration support |
| New Zealand Meteorological Services | Wellington, New Zealand | Technical Partner | Radar training, calibration and support; Technical training and support aviation, upper-air; forecasting training and support. |
| BoM (Australian Bureau of Meteorology) | Melbourne, Australia | Technical Partner | Technical training, radar and ICT expertise, bilateral NMHS support (beyond RIC/RTC functions) |
| JMA (Japan Meteorological Agency) | Tokyo/Tsukuba, Japan | Technical Partner | Advanced training, satellite data provision, calibration standards; support through JICA and WMO programmes |
| SPC (Pacific Community) | Noumea, New Caledonia & Suva, Fiji | Regional Organisation | Water resources, GIS, ICT, geoscience capacity development, hydrology training |
| WMO (World Meteorological Organization) Regional Offices | Singapore & Apia, Samoa | Regional Organisation | Coordinating regional meteorological activities, strengthening capacity development, and facilitating implementation of WMO programmes. Link between countries and WMO headquarters to improve service delivery, data exchange, and regional cooperation. |