# Procurement ToolKit – Annex B

Part of Weather Ready Pacific Standard, Optimisation & Maintenance Guidelines Document

Module: Automatic Weather Station (AWS) & Land (Optional)

Version: 1.0 | March 2026

When to Use This Module

Use this template when procuring:

- Complete AWS

- Replacement sensors for existing AWS

- Upgrade of telemetry or power systems

- Additional stations for network expansion

IMPORTANT NOTICE

This Procurement Toolkit is a guidance document only, developed under the Weather Ready Pacific (WRP) Programme to support standardised procurement of Earth observation equipment across Pacific Island Countries and Territories.

It is not a complete tender document. Users may need to adapt all content to:

- National procurement laws and policies

- Local environmental and workplace safety regulations

- Project-specific requirements and site conditions

- Applicable customs, import, and tax requirements

Professional judgment is required. Users should seek appropriate technical and legal advice before finalising any tender.

Limitation of Liability: The WRP Programme, its implementing partners, and supporting organisations accept no liability for any decisions made or actions taken based on this guidance.

Living Document: This toolkit will be updated periodically. Verify you have the latest version through SPREP or the WRP Programme.

Contact: Weather Ready Pacific Programme, SPREP, Apia, Samoa [include the email address of the WRP programme?]

## System Definition

Project Name: _________________________________

Number of station(s): _____

Location(s): _________________________________

Site types: ☐ Coastal ☐ Inland ☐ Mountain ☐ Urban ☐ Airport

Power available at site: ☐ Grid ☐ Solar only ☐ Generator ☐ Unknown

Cellular coverage: ☐ Yes ☐ No ☐ Partial ☐ Unknown

WIGOS Station Type: ☐ GBON ☐ RBON ☐ National Only

## Tier Selection

| Tier | Typical Use Case | WMO Classification | Recommended For |
| --- | --- | --- | --- |
| 1 (Reference) | Climate monitoring, research, GBON | Class 1 | National climate reference, capital cities, research sites |
| 2 (Synoptic) | Operational forecasting, aviation, RBON | Class 2 | Airport stations, forecast offices, major towns |
| 3 (Basic) | Sectoral applications, community monitoring | Class 3-4 | Agriculture, hydrology, schools, remote communities |

Selected Tier: ☐ 1 (Reference) ☐ 2 (Synoptic) ☐ 3 (Basic) ☐ Mixed (specify: _________)

Note: If a single site will contain equipment from different tiers (e.g., a Tier 1 wind sensor with a Tier 3 rain gauge), this must be clearly specified. The procurement documents should list the required tier for each sensor or system component.

** Your selected tier will dictate the rest of the tender

## System Architecture

This section defines how the system is built, how components interact, and how data flows from sensor to user.

  - Data Acquisition & Processing (Datalogger)

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Processor | 32-bit, industrial grade, fanless, -40°C to +70°C operating range | 16-bit or 32-bit, industrial grade, -30°C to +60°C | 16-bit, commercial grade, 0°C to +50°C |  |
| Memory | ≥128 MB, ≥1 year data storage at 1-minute intervals | ≥32 MB, ≥6 months data storage | ≥8 MB, ≥3 months data storage |  |
| Analog Channels | ≥12 differential or 24 single-ended, software configurable | ≥8 differential or 16 single-ended | ≥4 differential or 8 single-ended |  |
| Digital Channels | ≥8 digital I/O, ≥4 serial/RS232, SDI-12 (≥5 sensors) | ≥4 digital I/O, ≥2 serial, SDI-12 (≥3 sensors) | ≥2 digital I/O, SDI-12 capable |  |
| Scan Rate | ≤1 second for all channels | ≤5 seconds | ≤10 seconds |  |
| A/D Resolution | 24-bit | 16-bit | 16-bit |  |
| Accuracy | ±0.02% of reading | ±0.1% of reading | ±0.25% of reading |  |
| Time Synchronization | GPS-integrated, automatic daily sync, <1ms accuracy | GPS or NTP, <1 second accuracy | Manual or NTP when available |  |
| Real-Time Clock | Battery-backed, <2 minutes/year drift | Battery-backed, <5 minutes/year drift | Battery-backed, <10 minutes/year drift |  |
| Communications Ports | Ethernet, USB, RS-232, RS-485, maintenance port | USB, RS-232, maintenance port | USB or serial, maintenance port |  |
| Data Formats | Native: binary; Export: CSV, JSON, BUFR-capable | Native: binary; Export: CSV, JSON | Export: CSV |  |
| Remote Programming | Full remote configuration and firmware update | Remote configuration | Manual only |  |
| Diagnostics | Comprehensive self-diagnostics, watchdog timer, voltage monitoring | Basic self-diagnostics, watchdog timer | Watchdog timer |  |

Additional Requirements for Data Acquisition for all tiers:

- Programmability for custom processing (averages, maxima, minima)

- Configurable alarms for sensor limits, power failure, communication loss

- Automatic data retention during power loss

- Maintenance port for local laptop connection

- Enclosure: IP66 minimum if not in separate shelter

  - Sensor Complement

| Sensor Type | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Air Temperature | ✓ Mandatory | ✓ Mandatory | ✓ Mandatory | ☐ |
| Relative Humidity | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Atmospheric Pressure | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Wind Speed | ✓ Mandatory | ✓ Mandatory | ✓ Mandatory | ☐ |
| Wind Direction | ✓ Mandatory | ✓ Mandatory | ✓ Mandatory | ☐ |
| Precipitation | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Solar Radiation | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Sunshine Duration | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Soil Temperature | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Soil Moisture | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Evaporation | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Present Weather | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Visibility | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |

Selected Sensors for This Procurement:

List the specific sensors required:

| Sensor | Quantity | Manufacturer Preference (if any) | Notes |
| --- | --- | --- | --- |

  - Data Transmission & Telemetry

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Communication | Dedicated: ☐ Cellular (4G/LTE) ☐ Satellite (BGAN/Starlink/VSAT) ☐ Radio | ☐ Cellular ☐ Satellite ☐ Radio (choose one) | ☐ Cellular ☐ Satellite ☐ Manual download |  |
| Secondary Communication | ☐ Cellular backup (if primary satellite) ☐ Satellite backup (if primary cellular) ☐ Iridium SBD | ☐ Cellular where available ☐ Not required | Not required |  |
| Data Protocol | TCP/IP, HTTP/HTTPS, MQTT, FTP/SFTP, push to NMHS server | TCP/IP, HTTP/HTTPS, FTP, push to NMHS server | FTP or email, or manual download |  |
| Transmission Frequency | Real-time: ≤5 minute intervals; Synoptic: hourly on the hour | ≤15 minute intervals; Synoptic hourly | Daily summary or on request |  |
| Data Formats | BUFR (mandatory for GBON), CSV, JSON; SYNOP generation | CSV, JSON; SYNOP generation preferred | CSV only |  |
| Data Plan | Minimum 3-year prepaid service included (for satellite) | Minimum 2-year prepaid service included | NMHS provides SIM or manual |  |
| Redundancy | Automatic failover between primary/secondary, store-and-forward during outages | Manual switchover if backup exists; store-and-forward | Store-and-forward only |  |
| Remote Monitoring | Full link quality monitoring, signal strength, data usage, alarms | Basic link monitoring | Not required |  |

Data Flow Diagram Requirement regardless of tier:

The bidder must provide a detailed data flow diagram showing:

- Path from sensor → datalogger → transmission → NMHS server / CliDE

- All data formats at each stage

- Backup paths and failover mechanisms

- Store-and-forward operation during outages

- Points of manual intervention

Integration Testing:

The contractor shall demonstrate successful data transmission to:

☐ NMHS central server / CliDE

☐ WIS 2.0 node or GTS connection

☐ Forecaster workstation display

  - Power System

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Power Source | ☐ Grid with UPS ☐ Solar (specify below) | ☐ Grid with battery ☐ Solar | ☐ Solar ☐ Battery only |  |
| Solar Panel | Minimum 200W, monocrystalline, marine-grade anodized frame, 25-year lifespan | Minimum 150W, monocrystalline, corrosion-resistant frame | Minimum 100W, polycrystalline |  |
| Charge Controller | MPPT, 30A minimum, temperature compensation, remote monitoring capable | MPPT or PWM, 20A, temperature compensation | PWM, 10A |  |
| Battery Bank | Deep-cycle AGM or Lithium, 200Ah minimum, 5-year design life, vented enclosure | Deep-cycle AGM, 150Ah minimum, 3-year design life | Deep-cycle lead-acid, 100Ah minimum |  |
| Autonomy | Minimum 14 days without sun, full operation | Minimum 7 days without sun, full operation | Minimum 3 days without sun, critical sensors only |  |
| Power Distribution | Dedicated distribution panel, circuit breakers, surge protection on all lines | Fused distribution, surge protection on main line | Fused distribution |  |
| Voltage Regulation | Regulated outputs: 12V, 5V; all sensors protected | Regulated 12V output | Battery voltage direct |  |
| Power Monitoring | Remote monitoring: voltage, current, battery health, charge status, alarms | Local display of voltage and charge status | Visual battery check |  |
| Enclosure | Lockable, weatherproof (IP66), ventilated, corrosion-resistant | Lockable, weatherproof (IP65), ventilated | Weatherproof, basic lock |  |

Solar Power Sizing Calculation:

The bidder must provide detailed power budget showing:

- System total power consumption (watts)

- Daily energy requirement (watt-hours)

- Solar panel sizing calculation

- Battery bank sizing calculation

- Expected autonomy under worst-case conditions

Note: 14-day autonomy is standard for AWS sites. For locations with extended cloudy periods (e.g., during wet season), larger systems may be required.

  - Physical Infrastructure

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Mast / Tower | 10m galvanized steel, tilt-down for maintenance, hot-dip galvanized after fabrication, corrosion-resistant coating | 10m galvanized steel, fixed or tilt-down, galvanized | 5-6m galvanized steel or aluminium, fixed |  |
| Mast Foundation | Engineered concrete foundation, certified for 250 km/h winds, site-specific design | Concrete foundation for 200 km/h winds | Ground screw or concrete base for 150 km/h winds |  |
| Instrument Crossarm | Stainless steel, adjustable, 2m length, sensor mounts at standard heights | Stainless steel or galvanized, fixed, 1.5m length | Galvanized, fixed, 1m length |  |
| Instrument Enclosure | Lockable fiberglass or stainless steel, IP66, sun shield, passive ventilation, internal light | Lockable fiberglass or powder-coated steel, IP65, sun shield | Lockable plastic or metal, weatherproof |  |
| Equipment Shelter | 10m x 10m fenced compound, 2m high chain-link with barbed wire, locked gate | 5m x 5m fenced compound, locked gate | Small enclosure or fenced area, lockable |  |
| Lightning Protection | Full system: lightning rod on mast, down conductor, ground rods (<10 ohm), surge arrestors on all lines | Lightning rod on mast, down conductor, ground rod, surge arrestors on power | Basic lightning rod or disconnect during storms |  |
| Grounding | Separate ground for power and signal, <10-ohm resistance, annual test point | Combined ground, <25-ohm resistance | Basic ground rod |  |
| Cabling | UV-resistant, rodent-proof, armoured where exposed, all cables in conduit with drip loops, labelled both ends | UV-resistant, protected in conduit, labelled | Outdoor cable, secured, labelled |  |
| Cable Entry | Watertight glands, drip loops, no water entry path | Watertight glands, drip loops | Basic cable entry |  |

Site Preparation Requirements:

The contractor shall provide detailed site preparation guide including:

- Foundation drawings and specifications

- Concrete requirements (volume, strength)

- Cable trenching depths and routing

- Grounding system installation details

- Site clearing and grading requirements

- Local regulatory approvals needed

- Safety clearances from obstacles

- Photographic documentation requirements

  - Software & Data Management

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Logger Configuration Software | Full-featured, graphical interface, offline programming capability, data retrieval, real-time monitoring | Full-featured, graphical interface, offline programming | Basic configuration tool |  |
| Data Viewing Software | Real-time and historical data display, graphing, export, multi-station dashboard | Real-time display, basic graphing | Not required |  |
| Data Management System | Compatible with CliDE, automatic data push, QC flagging, metadata management | Compatible with CliDE, data push | CSV export for manual import |  |
| WIS 2.0 Compliance | BUFR generation capability, push to WIS 2.0 node | BUFR generation preferred | Not required |  |
| SYNOP Generation | Automatic SYNOP in TAC and BUFR, configurable for GTS | SYNOP generation capability | Not required |  |
| Mobile Application | Optional: mobile app for data viewing and alerts | Optional | Not required |  |
| License Model | Perpetual license, no recurring fees | Perpetual license | Included |  |
| Software Updates | 5 years free updates included | 3 years updates included | As available |  |
| API Access | Full open API for integration with other systems | Documented data access methods | Not required |  |

  - Integration with Existing Systems

The proposed system must integrate with the following systems:

| System | Integration Requirement | Compliant? |
| --- | --- | --- |
| [e.g., CliDE] | Automatic data push in required format, real-time or scheduled |  |
| [e.g., WIS 2.0 Node] | BUFR generation, automatic transfer |  |
| [e.g., Forecaster Workstation] | Data available in real-time for display |  |
| [e.g., National Database] | Data format compatible with import procedures |  |
| [e.g., Telemetry System] | Compatible with existing telemetry server |  |

Integration Testing:

The contractor shall demonstrate successful integration during:

- Factory Acceptance Test (simulated)

- Site Acceptance Test (live)

## Technical Specification by Tier

The specifications in this table are derived from the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 5.2.1 (Tier Sensor Specification by Variable – Atmosphere (Near Surface)). Refer to the Guide for detailed sensor accuracy and uncertainty data.

  - Atmosphere (Near Surface) Sensor Specifications

Tier 1 - Reference Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Air Temperature | Accuracy: ±0.2°C, Resolution: 0.1°C, Range: -40°C to +60°C, in radiation shield, 1.25-2m height |  |  |
| Relative Humidity | Accuracy: ±2% RH, Resolution: 1%, Range: 0-100%, stable in high humidity |  |  |
| Atmospheric Pressure | Accuracy: ±0.2 hPa, Resolution: 0.1 hPa, Range: 500-1080 hPa, temperature compensated |  |  |
| Wind Speed | Accuracy: ±0.5 m/s (<5 m/s) or ±10% (>5 m/s), Resolution: 0.1 m/s, Range: 0-75 m/s, 10m height |  |  |
| Wind Gust | 3-second average, same accuracy as wind speed, sampling ≤3 seconds |  |  |
| Wind Direction | Accuracy: ±5°, Resolution: 1°, Range: 0-360°, referenced to True North |  |  |
| Precipitation | Accuracy: ±1 mm or 2%, Resolution: 0.1 mm, Range: 0-1000 mm/day, 300 cm² collector minimum, wind shield |  |  |
| Precipitation Intensity | Accuracy: ±0.2 mm/h or 5%, Resolution: 0.1 mm/h, Range: 0-500 mm/h |  |  |
| Solar Radiation | Accuracy: ±5% daily total, Resolution: 1 W/m², Range: 0-2000 W/m², ISO 9060 secondary standard |  |  |
| Sunshine Duration | Accuracy: ±0.1 h, Threshold: 120 W/m² direct radiation |  |  |

Special Requirements for Tier 1:

- All sensors must have factory calibration certificates traceable to ISO 17025

- Radiation shield must be naturally ventilated, multiple-plate design

- Precipitation gauge must include alter-type wind shield

- Dual sensor redundancy for critical variables (T, RH, P) recommended

Tier 2 - Synoptic Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Air Temperature | Accuracy: ±0.5°C, Resolution: 0.1°C, Range: -20°C to +60°C, in radiation shield |  |  |
| Relative Humidity | Accuracy: ±3% RH, Resolution: 1%, Range: 0-100% |  |  |
| Atmospheric Pressure | Accuracy: ±0.5 hPa, Resolution: 0.1 hPa, Range: 500-1080 hPa |  |  |
| Wind Speed | Accuracy: ±1 m/s (<5 m/s) or ±15% (>5 m/s), Resolution: 0.1 m/s, Range: 0-75 m/s |  |  |
| Wind Gust | 3-second average, sampling ≤5 seconds |  |  |
| Wind Direction | Accuracy: ±10°, Resolution: 2°, Range: 0-360°, True North |  |  |
| Precipitation | Accuracy: ±3 mm or 5%, Resolution: 0.2 mm, Range: 0-1000 mm/day, 200 cm² collector minimum |  |  |
| Precipitation Intensity | Accuracy: ±0.5 mm/h or 10%, Resolution: 0.2 mm/h, Range: 0-300 mm/h |  |  |
| Solar Radiation | Accuracy: ±10% daily total, Resolution: 1 W/m², Range: 0-2000 W/m², ISO 9060 first class |  |  |

Tier 3 - Basic Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Air Temperature | Accuracy: ±1.0°C, Resolution: 0.5°C, Range: -20°C to +60°C, in simple radiation shield |  |  |
| Relative Humidity | Accuracy: ±10% RH, Resolution: 5%, Range: 0-100% (optional) |  |  |
| Atmospheric Pressure | Accuracy: ±2.0 hPa, Resolution: 0.5 hPa, Range: 500-1080 hPa (optional) |  |  |
| Wind Speed | Accuracy: ±2.5 m/s (<5 m/s) or ±20% (>5 m/s), Resolution: 0.5 m/s, Range: 0-60 m/s |  |  |
| Wind Gust | 3-second average, sampling ≤10 seconds |  |  |
| Wind Direction | Accuracy: ±15°, Resolution: 5°, Range: 0-360° |  |  |
| Precipitation | Accuracy: ±5 mm or 10%, Resolution: 0.5 mm, Range: 0-500 mm/day, tipping bucket |  |  |

  - Land Sensor Specifications (Optional)

| Parameter | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Included? |
| --- | --- | --- | --- | --- |
| Soil Temperature | Accuracy: ±0.3°C, Resolution: 0.1°C, Range: 0°C to +55°C, depths: 5,10,20,50,100 cm | Accuracy: ±0.5°C, Resolution: 0.1°C, depths: 10,20,50 cm | Accuracy: ±1.0°C, Resolution: 0.5°C, single depth | ☐ |
| Soil Moisture | Accuracy: ±3% VWC, Resolution: 0.1%, depths: 10,20,50 cm | Accuracy: ±5% VWC, Resolution: 0.1%, depths: 10,20 cm | Accuracy: ±10% VWC, single depth | ☐ |
| Evaporation | Class A pan with automatic sensor or accurate evaporation sensor | Class A pan with manual measurement | Not required | ☐ |

## Spare Parts Requirements by Tier

#### Tier 1 Spare Parts Kit (per 5 stations)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Air Temperature/Relative Humidity sensor | 1 complete set |  |  |
|  | Barometric pressure sensor | 1 |  |  |
|  | Wind speed/direction sensor | 1 |  |  |
|  | Precipitation gauge (complete) | 1 |  |  |
|  | Solar radiation sensor | 1 (if installed) |  |  |
| Data Acquisition | Data logger (complete, programmed) | 1 |  |  |
|  | Logger power supply | 1 |  |  |
|  | Communication modem (same as installed) | 1 |  |  |
| Power System | Solar panel (same rating) | 1 |  |  |
|  | Charge controller | 1 |  |  |
|  | Battery (deep-cycle, same specification) | 2 |  |  |
| Infrastructure | Sensor cables - full set | 2 sets |  |  |
|  | Connectors and glands kit | 2 sets |  |  |
|  | Fuses, surge protectors | 2 sets |  |  |
|  | Mounting hardware for sensors | 1 set |  |  |
| Consumables | Desiccant packs | 10 |  |  |
|  | Cable ties (UV-stabilized) | 100 |  |  |
|  | Cleaning kit | 1 |  |  |
| Tools | Field calibration kit (T/RH reference, barometer reference) | 1 |  |  |
|  | Laptop with all software (optional) | 1 |  |  |
| TOTAL (Tier 1) |  |  |  |  |

#### Tier 2 Spare Parts Kit (per 5 stations)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Air Temperature/Relative Humidity sensor | 1 |  |  |
|  | Barometric pressure sensor | 1 |  |  |
|  | Wind speed/direction sensor | 1 |  |  |
|  | Precipitation gauge (tipping mechanism) | 1 set |  |  |
| Data Acquisition | Data logger (complete) | 1 |  |  |
|  | Communication modem | 1 |  |  |
| Power System | Battery (same specification) | 2 |  |  |
|  | Charge controller | 1 |  |  |
| Infrastructure | Critical cables | 1 set |  |  |
|  | Connectors and fuses | 1 set |  |  |
| TOTAL (Tier 2) |  |  |  |  |

#### Tier 3 Spare Parts Kit (per 10 stations)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Air Temperature sensor | 1 |  |  |
|  | Wind sensor (if installed) | 1 |  |  |
|  | Rain gauge bucket/tipping mechanism | 1 |  |  |
| Power | Battery | 2 |  |  |
| Infrastructure | Basic toolkit | 1 |  |  |
| TOTAL (Tier 3) |  |  |  |  |

## Maintenance Schedule

Tier 1 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Monthly | Visual inspection of all components, clean radiation shield, check for corrosion, check cable condition, check enclosure seals, verify data transmission, review data quality | 2 hours | Technician |
| Quarterly | Full sensor inspection, battery terminal cleaning and torque check, solar panel cleaning, guy wire tension check, foundation inspection, complete system diagnostics, data offload verification | 4 hours | Technician |
| Bi-annually | Field verification with portable reference (T/RH/P), rain gauge calibration check, anemometer bearing check, lightning protection system check | 1 day | Specialist |
| Annually | Full calibration of all sensors (return to lab or exchange), mast inspection for corrosion, repaint if needed, battery capacity test, grounding resistance test, software/firmware updates | 2-3 days | Specialist + Manufacturer |
| 5-Year | Major refurbishment: replace all sensors, replace batteries, replace critical cables, full system recertification | 1 week | Manufacturer |

Tier 2 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Monthly (remote) | Remote data check | 30 min | Technician |
| Quarterly | Visual inspection, clean sensors, check cables, battery check, verify data transmission | 1-2 hours | Technician |
| Bi-annually | Detailed inspection, field verification of critical sensors (T/RH/P), rain gauge cleaning, solar panel cleaning | 4 hours | Technician |
| Annually | Full system check, battery replacement if needed, calibration verification, software updates | 1 day | Technician + Support |

Tier 3 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Bi-annually | Visual inspection, clean sensors, battery check, verify data | 1 hour | Technician or Community Observer |
| Annually | Full check, battery replacement as needed | 2-3 hours | Technician |

## Field Verification & Calibration

This section defines requirements for two distinct activities. Field Verification is the on-site check of instrument performance against a portable reference. Calibration is a formal, traceable comparison against a national or international standard, typically performed in a laboratory setting. See the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 7.4 for more detail.

  - Field Verification Equipment

| Item | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Temperature Reference | Accuracy ±0.2°C, NIST traceable | Accuracy ±0.5°C | Not required | ☐ |
| Humidity Reference | Accuracy ±2% RH | Accuracy ±5% RH | Not required | ☐ |
| Barometric Reference | Accuracy ±0.2 hPa, calibrated | Accuracy ±0.5 hPa | Not required | ☐ |
| Rain Gauge Calibrator | Graduated cylinder with stand, drip rate control | Graduated cylinder | Not required | ☐ |
| Anemometer Test Kit | Portable wind source, tachometer | Basic check | Not required | ☐ |

  - Calibration Requirements

| Component | Tier 1 | Tier 2 | Tier 3 |
| --- | --- | --- | --- |
| Temperature Sensor | Annual laboratory calibration | 2-year calibration | As-failed replacement |
| Humidity Sensor | Annual laboratory calibration | 2-year calibration | As-failed replacement |
| Pressure Sensor | Annual laboratory calibration | 2-year calibration | As-failed replacement |
| Wind Sensors | 2-year verification | 3-year verification | As failed |
| Rain Gauge | Annual field verification, 2-year lab | Annual field verification | As failed |
| Field Verification Equipment | Annual calibration at accredited lab | Annual calibration | N/A |

## Warranty & Services

  - Warranty Requirements

| Component | Tier 1 | Tier 2 | Tier 3 | Bidder Offer |
| --- | --- | --- | --- | --- |
| Warranty Period | Minimum 3 years comprehensive | Minimum 2 years comprehensive | Minimum 1 year comprehensive |  |
| Coverage | Parts, labor, shipping both ways | Parts, labor, shipping to vendor | Parts only, return to vendor |  |
| Response Time | <24 hours for critical failure acknowledgment | <48 hours for critical failure acknowledgment | <5 working days |  |
| Remote Support | 24/7 hotline, remote diagnostics included | Business hours hotline (8am-5pm local), remote diagnostics | Email support |  |
| On-site Support | Guaranteed on-site within 7 days for critical failures (travel conditions permitting) | On-site within 14 days if remote cannot resolve | Not included |  |
| Parts Replacement | Advance replacement shipped within 48 hours | Replacement shipped within 1 week | Replacement shipped within 1 month |  |

  - Service Level Agreement (SLA) Requirements

The contractor shall provide a detailed SLA including:

| SLA Element | Tier 1 Requirement | Tier 2 Requirement | Bidder Offer |
| --- | --- | --- | --- |
| Availability target | 99% data availability (excluding scheduled maintenance) | 95% data availability |  |
| Scheduled maintenance | Maximum 3 days per year, advance notice 4 weeks | Maximum 4 days per year, advance notice 2 weeks |  |
| Helpdesk hours | 24/7/365 | Business hours (8am-5pm local time) with after-hours emergency contact |  |
| Critical failure response | <2 hours acknowledgment, <24 hours remote attempt | <24 hours acknowledgment |  |
| Spare parts delivery | <1 week to site (within region) | <2 weeks to site |  |
| Software updates | All updates and security patches included for warranty period | Critical security patches included |  |
| Performance reporting | Monthly report on system performance, issues, resolutions | Quarterly report |  |

  - Post-Warranty Support Options

The bidder must provide pricing for post-warranty support:

| Option | Description | Annual Cost | Notes |
| --- | --- | --- | --- |
| Option 1: Comprehensive | Full coverage as during warranty |  |  |
| Option 2: Parts only | Spare parts at fixed prices, no labour |  |  |
| Option 3: Time & materials | Support charged as needed |  |  |
| Option 4: Training refresh | Refresher training for new staff |  |  |

## Training Requirements

| Training Type | Audience | Duration | When | Must Include |
| --- | --- | --- | --- | --- |
| Operator training | Daily users, data monitors | 2-3 days | Before or during installation | Data viewing, basic troubleshooting, when to call for help, data quality checks |
| Maintenance training | Technical staff, technicians | 5 days | During installation (hands-on) | System architecture, sensor installation, preventive maintenance, diagnostics, repairs, spare parts replacement |
| Calibration training | Senior technicians | 2-3 days | After installation | Field verification procedures, use of calibration kit, understanding calibration certificates |
| Software training | Data managers, IT staff | 2-3 days | After installation | Logger configuration, data retrieval, software use, integration with NMHS systems |
| Installation training | Technical staff | Part of maintenance | During installation | Hands-on installation of at least one complete station |
| Train-the-trainer | Senior technicians | 2 days | Optional | How to train new staff, create local procedures |

Training materials required:

☐ User manuals (printed + digital) in English

☐ Maintenance manuals with exploded diagrams

☐ Quick reference guides (laminated for field use) – 2 copies per station

☐ Video tutorials (recommended)

☐ Spare parts catalog with part numbers and photos

☐ Troubleshooting flowchart (laminated)

☐ Preventative maintenance checklist (laminated)

Training Venue:

- First session (operator): At NMHS or online

- Second session (maintenance/installation): At NMHS, including field installation at 1-2 sites

- Software training: At NMHS computer lab

Training Evaluation:

- Pre- and post-training assessments required

- Practical demonstration of competency (each participant must perform tasks)

- Certificate of completion for each participant

- Training report with recommendations for additional support

## Technical Documentation

The contractor must provide comprehensive technical documentation:

  - System Documentation

| Document | Description | Format | Quantity |
| --- | --- | --- | --- |
| System Design Document | Overall architecture, component specifications, interface definitions, wiring diagrams | Digital (PDF) + 2 hard copies | 2 sets |
| Installation Manual | Step-by-step installation instructions, site preparation requirements, commissioning procedures, alignment specifications | Digital (PDF) + 2 hard copies | 2 sets |
| User Manual | Daily operation procedures, data viewing, basic troubleshooting, maintenance alerts | Digital (PDF) + 2 hard copies | 3 sets |
| Maintenance Manual | Preventive maintenance schedules, troubleshooting guides, repair procedures, part replacement instructions | Digital (PDF) + 2 hard copies | 3 sets |
| Technical Drawings | Mechanical drawings, electrical schematics, cabling diagrams, mast foundation details, sensor mounting details | Digital (PDF + CAD if available) + 2 hard copies | 2 sets |
| Parts Catalog | Complete list of all components with part numbers, suppliers, cross-reference, exploded views | Digital (PDF) + 2 hard copies | 3 sets |
| Software Documentation | User guides, administrator guides, API documentation, data format specifications, configuration guides | Digital (PDF) | 3 sets |
| Calibration Procedures | Step-by-step calibration instructions, traceability information, calibration interval recommendations | Digital (PDF) + 2 hard copies | 3 sets |
| Safety Manual | All safety procedures, hazard identification, emergency response, safe work practices | Digital (PDF) + 2 hard copies | 3 sets |
| Warranty Documentation | Warranty terms, registration, claim procedures, service contact information | Digital (PDF) + 2 hard copies | 3 sets |
| As-Built Documentation | Final as-installed drawings, configuration settings, test results, photographs (within 30 days of completion) | Digital (PDF) + 2 hard copies | 2 sets |

  - Documentation Quality Requirements

- All documentation must be in English

- Language must be clear, avoiding unnecessary jargon

- Diagrams and illustrations must be used extensively

- Troubleshooting guides must include decision trees

- All documents must include revision control and date

- Hard copies must be durable (weather-resistant covers for field manuals)

- Digital copies must be searchable PDFs

- Critical procedures must have bilingual summaries (English + local language) where possible

  - Site-Specific Documentation

For each installed station, the contractor shall provide:

- Station location map and coordinates (WGS84)

- Site photographs (4 directions from mast, close-ups of each component)

- Sensor heights and orientations

- Obstruction diagram (for wind sensors)

- Grounding resistance measurement

- Commissioning test results

- Configuration file for datalogger

- Network and communication settings

- Contact list for support

## Lifecycle Cost Calculation

| Cost Category |  | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | 5-Year Total |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Capital costs | Equipment (sensors, logger, etc.) |  |  |  |  |  |  |
|  | Mast and infrastructure |  |  |  |  |  |  |
|  | Power system |  |  |  |  |  |  |
|  | Telemetry equipment |  |  |  |  |  |  |
|  | Installation |  |  |  |  |  |  |
|  | Training |  |  |  |  |  |  |
| Recurring costs | Data plans (cellular/satellite) |  |  |  |  |  |  |
|  | Consumables (desiccant, etc.) |  |  |  |  |  |  |
|  | Batteries (if not included in capital) |  |  |  |  |  |  |
| Maintenance | Spare parts (planned) |  |  |  |  |  |  |
|  | Calibration services |  |  |  |  |  |  |
|  | Software licenses / support |  |  |  |  |  |  |
|  | Support contract (post-warranty) |  |  |  |  |  |  |
|  | Travel for maintenance |  |  |  |  |  |  |
| TOTAL |  |  |  |  |  |  |  |

Notes:

- Include all costs in [USD or local currency]

- Assume normal operation, no catastrophic failures

- Include inflation if multi-year contract

- Separate optional items clearly

## Pacific Requirements

All AWS systems must meet these additional requirements:

| Requirement | Specification | Compliant? |
| --- | --- | --- |
| Corrosion protection | Marine-grade materials (316 stainless steel, marine aluminum, or powder-coated galvanized steel), sealed connectors, tropical coating on all outdoor components |  |
| UV resistance | All outdoor components UV-stabilized (cabling, enclosures, plastic parts) - minimum 5-year UV life |  |
| Cyclone resistance | Mast and sensors rated to 250 km/h (Tier 1/2), 180 km/h (Tier 3); all components securely fastened |  |
| Salt spray protection | All outdoor components sealed to IP66 minimum, marine-grade fasteners (stainless steel), no exposed mild steel |  |
| Waterproofing | IP66 for all outdoor enclosures, IP67 for connectors, sealed cable entries with glands and drip loops |  |
| Humidity tolerance | 0-100% condensing environment, all electronics conformally coated or protected |  |
| Temperature range | Operational: 0°C to +50°C (typical), survival: -10°C to +60°C |  |
| Mould/fungus resistance | All materials treated or selected for tropical resistance, no organic materials in enclosures |  |
| Insect/pest protection | Screened ventilation openings, sealed cable entries, ant-proof considerations |  |
| Lightning protection | Complete system protection with redundancy, compliance with IEC 62305 or equivalent, <10-ohm ground |  |
| Power autonomy | Minimum 14 days without sun (Tier 1), 7 days (Tier 2), 3 days (Tier 3) - full operation |  |
| Solar panel orientation | Adjustable mounts for optimal angle at latitude (typically 0-15° in tropics) |  |
| Data retrieval during outage | Store-and-forward capability for minimum period equal to power autonomy |  |
| Local supportability | Common connectors, standard tools for maintenance, no proprietary fasteners |  |

## Compliance Checklist

Before finalising your tender, verify:

| Check | Completed? |
| --- | --- |
| Tier selection matches your operational needs | ☐ |
| All required variables selected | ☐ |
| System architecture fully defined | ☐ |
| All sensor specifications are included and match the tier selected | ☐ |
| Data transmission requirements specified and appropriate for the site | ☐ |
| Power system defined correctly for the location | ☐ |
| Physical infrastructure requirements clear and appropriate for cite conditions | ☐ |
| Software and integration requirements specified | ☐ |
| Spare parts kit matches tier requirement and network size | ☐ |
| Maintenance schedule defined and resourced | ☐ |
| Field verification equipment included | ☐ |
| Warranty and SLA requirements specified | ☐ |
| Training requirements fully defined and specified | ☐ |
| Technical documentation requirements included | ☐ |
| Pacific requirements included | ☐ |
| Lifecycle cost worksheet completed (5-year TCO) | ☐ |
| Budget confirmed for capital AND 5-year operation | ☐ |
| Site surveys completed for all locations | ☐ |

## Additional Information for Bidders

  - Site Information (to be completed by NMHS)

| Site | Latitude | Longitude | Elevation | Access | Power | Comms | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |

  - Existing Equipment at Sites

| Site | Existing Logger | Existing Sensors | Compatible? |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |

  - Special Local Requirements

Import duties: ☐ NMHS responsible ☐ Bidder responsible

Customs clearance: ☐ NMHS assists ☐ Bidder responsible

Local taxes: ☐ Applicable ☐ Exempt (specify)

Work permits: ☐ Required ☐ Not required

Site access restrictions: _________________________________