# WRP Procurement ToolKit – Annex B

Part of Weather Ready Pacific Standard, Optimisation & Maintenance Guidelines Document

Module: Freshwater & Hydrological Monitoring Systems

Version: 1.0 | March 2026

When to Use This Module

Use this template when procuring:

- River/stream level gauging stations

- Rainfall stations (hydrological network)

- Water quality monitoring stations

- Groundwater monitoring systems

- Flood warning sensors and telemetry

- Complete hydrological network expansion

IMPORTANT NOTICE

This Procurement Toolkit is a guidance document only, developed under the Weather Ready Pacific (WRP) Programme to support standardised procurement of Earth observation equipment across Pacific Island Countries and Territories.

It is not a complete tender document. Users may need to adapt all content to:

- National procurement laws and policies

- Local environmental and workplace safety regulations

- Project-specific requirements and site conditions

- Applicable customs, import, and tax requirements

Professional judgment is required. Users should seek appropriate technical and legal advice before finalising any tender.

Limitation of Liability: The WRP Programme, its implementing partners, and supporting organisations accept no liability for any decisions made or actions taken based on this guidance.

Living Document: This toolkit will be updated periodically. Verify you have the latest version through SPREP or the WRP Programme.

Contact: Weather Ready Pacific Programme, SPREP, Apia, Samoa [include the email address of the WRP programme?]

## System Definition

Project Name: _________________________________

Number of station(s): _____

Station types:

☐ River level only _____ stations

☐ River level + rainfall _____ stations

☐ River level + rainfall + water quality _____ stations

☐ Groundwater level _____ stations

☐ Rainfall only _____ stations

Location(s): _________________________________

River/catchment names: _________________________________

Primary purpose: ☐ Flood warning ☐ Water resources ☐ Climate monitoring ☐ Research

Power available at site: ☐ Grid ☐ Solar only ☐ No power available ☐ Unknown

Cellular coverage: ☐ Yes ☐ No ☐ Partial ☐ Unknown

Flood risk at site: ☐ Low ☐ Medium ☐ High (station may be inundated) ☐ Unknown

Site access: ☐ Road ☐ Boat ☐ Walking only ☐ Helicopter

## Tier Selection

| Tier | Typical Use Case | NEMS Equivalent | Recommended For |
| --- | --- | --- | --- |
| 1 (Reference) | Long-term climate hydrology, research, benchmark stations | QC600 | Major rivers, national reference, climate change monitoring, research catchments |
| 2 (Synoptic) | Operational flood warning, water resource management | QC500 | Flood-prone rivers, water supply catchments, major tributaries |
| 3 (Basic) | Community-based monitoring, remote catchments, education | QC400 | Small catchments, community flood warning, schools, remote sites |

Selected Tier: ☐ 1 (Reference) ☐ 2 (Synoptic) ☐ 3 (Basic) ☐ Mixed (specify: _________)

Note: If a single site will contain equipment from different tiers (e.g., a Tier 1 wind sensor with a Tier 3 rain gauge), this must be clearly specified. The procurement documents should list the required tier for each sensor or system component.

** Your selected tier will dictate the rest of the tender

## System Architecture

This section defines how the system is built, how components interact, and how data flows from sensor to user.

  - Data Acquisition & Processing (Datalogger)

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Processor | 32-bit, industrial grade, fanless, -40°C to +70°C operating range | 16-bit or 32-bit, industrial grade, -30°C to +60°C | 16-bit, commercial grade, 0°C to +50°C |  |
| Memory | ≥128 MB, ≥1 year data storage at 5-minute intervals (including high-flow events) | ≥32 MB, ≥6 months data storage | ≥8 MB, ≥3 months data storage |  |
| Analog Channels | ≥8 differential or 16 single-ended, software configurable | ≥4 differential or 8 single-ended | ≥2 differential or 4 single-ended |  |
| Digital Channels | ≥4 digital I/O, ≥2 serial/RS232, SDI-12 (≥5 sensors) | ≥2 digital I/O, ≥1 serial, SDI-12 capable | SDI-12 capable |  |
| Pulse Counters | ≥2 high-speed pulse counters (for rain gauge, shaft encoder) | ≥2 pulse counters | ≥1 pulse counter |  |
| Scan Rate | ≤1 second for critical sensors (level, rainfall) | ≤5 seconds | ≤15 seconds |  |
| A/D Resolution | 24-bit | 16-bit | 16-bit |  |
| Accuracy | ±0.02% of reading | ±0.1% of reading | ±0.25% of reading |  |
| Time Synchronization | GPS-integrated or NTP, automatic daily sync, <1ms accuracy | GPS or NTP, <1 second accuracy | Manual or NTP when available |  |
| Real-Time Clock | Battery-backed, <2 minutes/year drift | Battery-backed, <5 minutes/year drift | Battery-backed, <10 minutes/year drift |  |
| Event Recording | Ability to log at higher frequency during flood events (triggered by rate of rise) | Event-based logging capability | Standard interval logging only |  |
| Communications Ports | Ethernet, USB, RS-232, RS-485, maintenance port | USB, RS-232, maintenance port | USB or serial, maintenance port |  |
| Data Formats | Native: binary; Export: CSV, JSON, WaterML2 | Export: CSV, JSON | Export: CSV |  |
| Remote Programming | Full remote configuration and firmware update | Remote configuration | Manual only |  |
| Diagnostics | Comprehensive self-diagnostics, watchdog timer, voltage monitoring, sensor health checks | Basic self-diagnostics, watchdog timer | Watchdog timer |  |

Additional Datalogger Requirements for Hydrology:

- Rate-of-Rise Trigger: Must support event logging when water level rises faster than defined threshold (e.g., >0.5m/hour)

- Stage-Discharge Calculation: Ability to store rating curve and calculate discharge in real-time (Tier 1/B)

- Battery Voltage Logging: Must log battery voltage with each reading for power system monitoring

- Sensor Failure Detection: Automatic flagging of sensor out-of-range or failure conditions

  - Sensor Complement

| Sensor Type | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Water Level (River/Stream) | ✓ Mandatory | ✓ Mandatory | ✓ Mandatory | ☐ |
| Rainfall (Tipping Bucket) | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Water Temperature | ✓ Mandatory | ☐ Optional | ☐ Optional | ☐ |
| Water Quality - pH | ☐ Optional | ☐ Optional | ✓ Mandatory | ☐ |
| Water Quality - Conductivity | ☐ Optional | ☐ Optional | ✓ Mandatory | ☐ |
| Water Quality - Turbidity | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Water Quality – Dissolved Oxygen | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Barometric Pressure | ✓ Mandatory (for compensation) | ✓ Mandatory (for compensation) | ☐ Optional | ☐ |
| Air Temperature | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Groundwater Level | ✓ Mandatory if groundwater site | ✓ Mandatory if groundwater site | ✓ Mandatory if groundwater site | ☐ |

Selected Sensors for This Procurement:

List the specific sensors required:

| Sensor | Quantity | Manufacturer Preference (if any) | Notes (e.g., depth range, expected max level) |
| --- | --- | --- | --- |

  - Water Level Sensor Options

| Sensor Type | Advantages | Disadvantages | Suitable for Tier | Selected? |
| --- | --- | --- | --- | --- |
| Vented Pressure Transducer | High accuracy, stable, direct reading | Requires vented cable, desiccant maintenance, drift over time | 1, 2 | ☐ |
| Non-Vented (Absolute) Pressure with Barometric Compensation | Simpler cable, no vent tube | Requires separate barometric sensor, compensation calculation | 1, 2, 3 | ☐ |
| Radar Level Sensor | Non-contact, no moving parts, low maintenance, good for flood sites | Higher cost, requires stable mounting, affected by debris/foam | 1, 2 | ☐ |
| Ultrasonic Level Sensor | Non-contact, lower cost | Affected by temperature, wind, turbulence; less accurate | 2, 3 | ☐ |
| Shaft Encoder / Float & Pulley | Simple, reliable, direct mechanical measurement | Moving parts, requires stilling well, maintenance | 1, 2 | ☐ |
| Bubbler System | No sensor in water, good for silty environments | Compressed gas required, maintenance, slower response | 1, 2 | ☐ |
| Staff Gauge (Manual) | Zero cost, no power required | Manual reading only, no telemetry | All (as backup) | ☐ |

Selected Sensor Type: _________________________________

Justification: ________________________________

  - Data Transmission & Telemetry

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Communication | Dedicated: ☐ Cellular (4G/LTE) ☐ Satellite (BGAN/Starlink/VSAT) ☐ Radio | ☐ Cellular ☐ Satellite ☐ Radio (choose one) | ☐ Cellular ☐ Satellite ☐ Manual download |  |
| Secondary Communication | ☐ Cellular backup (if primary satellite) ☐ Satellite backup (if primary cellular) ☐ Iridium SBD | Recommended where feasible | Not required |  |
| Data Protocol | TCP/IP, HTTP/HTTPS, MQTT, FTP/SFTP, push to NMHS server | TCP/IP, HTTP/HTTPS, FTP, push to NMHS server | FTP or email, or manual download |  |
| Transmission Frequency | Real-time: ≤5-minute intervals during floods, ≤15 minutes normal; Event-based trigger for rapid rise | ≤15-minute intervals; Event-based trigger optional | Daily summary or on request |  |
| Flood Alert Transmission | Immediate transmission when flood threshold exceeded or rapid rise detected | Immediate transmission when threshold exceeded | Manual alert only |  |
| Data Formats | WaterML2, CSV, JSON; Compatible with national hydrology database | CSV, JSON | CSV only |  |
| Data Plan | Minimum 3-year prepaid service included (for satellite) | Minimum 2-year prepaid service included | NMHS provides SIM or manual |  |
| Redundancy | Automatic failover between primary/secondary, store-and-forward during outages | Manual switchover if backup exists; store-and-forward | Store-and-forward only |  |
| Remote Monitoring | Full link quality monitoring, signal strength, data usage, alarms | Basic link monitoring | Not required |  |

Flood Warning Integration:

System must be capable of sending alerts directly to:

- NMHS flood warning system

- SMS alerts to registered users

- Email alerts

- Siren activation (if applicable)

Data Flow Diagram Requirement regardless of tier:

The bidder must provide a detailed data flow diagram showing:

- Path from sensor → datalogger → transmission → NMHS server

- All data formats at each stage

- Alert paths for flood warnings

- Backup paths and failover mechanisms

- Store-and-forward operation during outages

- Points of manual intervention

Integration Testing:

The contractor shall demonstrate successful data transmission to:

☐ NMHS central server

☐ WIS 2.0 node or GTS connection

☐ Forecaster workstation display

  - Power System

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Power Source | ☐ Grid with UPS ☐ Solar (specify below) | ☐ Grid with battery ☐ Solar | ☐ Solar ☐ Battery only |  |
| Solar Panel | Minimum 100W (higher in shaded areas), monocrystalline, marine-grade anodized frame, 25-year lifespan | Minimum 80W, monocrystalline, corrosion-resistant frame | Minimum 50W, polycrystalline |  |
| Charge Controller | MPPT, 20A minimum, temperature compensation, remote monitoring capable | MPPT or PWM, 15A, temperature compensation | PWM, 10A |  |
| Battery Bank | Deep-cycle AGM or Lithium, 100Ah minimum, 5-year design life, vented enclosure | Deep-cycle AGM, 80Ah minimum, 3-year design life | Deep-cycle lead-acid, 50Ah minimum |  |
| Autonomy | Minimum 30 days without sun (critical for flood season reliability) | Minimum 14 days without sun | Minimum 7 days without sun |  |
| Power Distribution | Dedicated distribution panel, circuit breakers, surge protection on all lines | Fused distribution, surge protection on main line | Fused distribution |  |
| Voltage Regulation | Regulated outputs: 12V, 5V; all sensors protected | Regulated 12V output | Battery voltage direct |  |
| Power Monitoring | Remote monitoring: voltage, current, battery health, charge status, alarms | Local display of voltage | Visual battery check |  |
| Enclosure | Lockable, weatherproof (IP66), ventilated, corrosion-resistant | Lockable, weatherproof (IP65), ventilated | Weatherproof, basic lock |  |

Important Note for Hydrology Sites:

30-day autonomy is required because:

- Flood events occur during extended cloudy/rainy periods

- Sites may be inaccessible for weeks during floods

- Missing flood peaks due to power failure = complete system failure

The bidder must account for site-specific shading conditions (many hydrology sites are in shaded valleys) and provide larger panels where necessary.

Solar Power Sizing Calculation:

Many hydrology sites are in shaded valleys or forested areas. The bidder must:

- Account for site-specific shading conditions

- Provide larger panels where necessary

- Include battery sizing for extended cloudy periods during wet season (when flooding occurs)

  - Physical Infrastructure

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Sensor Mounting - River | Stilling well (if using pressure sensor) OR sturdy bridge/piers mount (if radar) OR dedicated sensor pole | Stilling well or bridge mount or sensor pole | Simple mount to existing structure |  |
| Stilling Well | 150-300mm diameter PVC or steel, perforated intake, stilling well cap, access for cleaning, de-silting provision | 150mm diameter, perforated, accessible | Not required if using non-contact sensor |  |
| Sensor Pole / Mast | 4-6m galvanized steel, concrete foundation, guyed if necessary, corrosion-resistant coating | 3-4m galvanized steel, concrete foundation | 2-3m galvanized steel, ground screw |  |
| Instrument Enclosure | Lockable fiberglass or stainless steel, IP66, sun shield, passive ventilation, internal light, anti-condensation measures | Lockable fiberglass or powder-coated steel, IP65, sun shield | Lockable plastic or metal, weatherproof |  |
| Fencing / Security | 5m x 5m fenced compound (if accessible to public), 2m high chain-link, locked gate | Fenced area or locked enclosure | Locked enclosure only |  |
| Lightning Protection | Full system: lightning rod on mast, down conductor, ground rods (<10 ohm), surge arrestors on all lines | Lightning rod on mast, down conductor, ground rod, surge arrestors on power | Basic lightning rod or disconnect during storms |  |
| Grounding | Separate ground for power and signal, <10 ohm resistance | Combined ground, <25 ohm resistance | Basic ground rod |  |
| Cabling | UV-resistant, rodent-proof, armoured where exposed, all cables in conduit with drip loops, labelled both ends | UV-resistant, protected in conduit, labelled | Outdoor cable, secured, labelled |  |
| Cable Entry | Watertight glands, drip loops, no water entry path | Watertight glands, drip loops | Basic cable entry |  |
| Staff Gauge | Permanent staff gauge installed at site for manual verification (1-5m length as needed) | Permanent staff gauge recommended | Portable staff gauge or marked reference |  |

Site Preparation Requirements:

The contractor shall provide detailed site preparation guide including:

- Stilling well installation details (depth, intake design, anti-silting measures)

- Foundation specifications for sensor pole

- Cable trenching requirements (depth for protection)

- Grounding system installation

- Site access requirements

- Flood-proofing measures

- Local regulatory approvals needed (river works permits)

  - Software & Data Management

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Logger Configuration Software | Full-featured, graphical interface, offline programming capability, data retrieval, real-time monitoring | Full-featured, graphical interface, offline programming | Basic configuration tool |  |
| Data Viewing Software | Real-time and historical data display, hydrograph generation, rating curve application, flood level display | Real-time display, basic hydrographs | Not required |  |
| Rating Curve Management | Ability to store and apply multiple rating curves, stage-discharge calculation, curve updates | Rating curve capability | Not required |  |
| Flood Alert Software | Configurable thresholds, multiple alert levels, SMS/email integration, siren control if applicable | Basic threshold alerts | Not required |  |
| Data Management System | Compatible with national hydrology, automatic data push | Compatible with national database, data push | CSV export for manual import |  |
| WaterML2 Compliance | Full WaterML2 output capability for international data exchange | WaterML2 preferred | Not required |  |
| Web Interface | Optional: web-based data viewing for public and emergency managers | Optional | Not required |  |
| Mobile Application | Optional: mobile app for data viewing and alerts | Optional | Not required |  |
| License Model | Perpetual license, no recurring fees | Perpetual license | Included |  |
| Software Updates | 5 years free updates included | 3 years updates included | As available |  |

  - Integration with Existing Systems

The proposed system must integrate with the following systems:

| System | Integration Requirement | Compliant? |
| --- | --- | --- |
| [e.g., National Hydrology Database] | Automatic data push in required format, real-time or scheduled |  |
| [e.g., Flood Warning System] | Real-time data feed with threshold alerts |  |
| [e.g., Web Display] | Data available for public website |  |
| [e.g., WIS 2.0] | WaterML2 output capability |  |

Integration Testing:

The contractor shall demonstrate successful integration during:

- Factory Acceptance Test (simulated)

- Site Acceptance Test (live at installation)

## Technical Specification by Tier

The specifications in this table are derived from the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 5.2.4 (Tier Sensor Specification by Variable – Freshwater & Hydrology). Refer to the Guide for detailed sensor accuracy and uncertainty data.

  - Water Level Sensor Specifications

Tier 1 - Reference Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Measurement Range | 0-10m minimum (or site-specific required range) |  |  |
| Accuracy | ±0.05% of full scale (e.g., ±5mm over 10m range) or ±3mm, whichever is greater |  |  |
| Resolution | 1 mm |  |  |
| Long-term Stability | <0.1% of full scale per year |  |  |
| Temperature Compensation | Full automatic compensation over 0°C to +40°C |  |  |
| Sampling Rate | Continuous, configurable from 1 second to 1 hour |  |  |
| Output | SDI-12, 4-20mA, or RS-232/485 |  |  |
| Over-range Protection | 2x rated range without damage |  |  |
| Materials | Marine-grade materials (titanium, 316 stainless steel, or ceramic for wetted parts) |  |  |
| Ingress Protection | IP68 (permanent immersion) |  |  |

Special Requirements for Tier 1:

- Dual sensor redundancy recommended for critical sites

- Factory calibration certificate traceable to ISO 17025 or national standard

- Desiccant system for vented sensors (replaceable)

- Lightning protection integrated

Tier 2 - Synoptic Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Measurement Range | 0-10m minimum (or site-specific) |  |  |
| Accuracy | ±0.1% of full scale (e.g., ±10mm over 10m range) |  |  |
| Resolution | 1-2 mm |  |  |
| Long-term Stability | <0.2% of full scale per year |  |  |
| Temperature Compensation | Automatic compensation |  |  |
| Sampling Rate | Configurable from 5 seconds to 1 hour |  |  |
| Output | SDI-12 or 4-20mA |  |  |
| Materials | Corrosion-resistant materials for freshwater |  |  |
| Ingress Protection | IP68 |  |  |

Tier 3 - Basic Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Measurement Range | 0-5m minimum |  |  |
| Accuracy | ±0.25% of full scale or ±1 cm |  |  |
| Resolution | 1 cm |  |  |
| Sampling Rate | Configurable from 1 minute to 1 hour |  |  |
| Output | SDI-12 or analog |  |  |
| Ingress Protection | IP67 |  |  |

  - Rainfall Sensor Specifications (Hydrological Network)

| Parameter | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Included? |
| --- | --- | --- | --- | --- |
| Type | Tipping bucket with wind shield | Tipping bucket | Tipping bucket or manual | ☐ |
| Collector Area | 400 cm² minimum | 200 cm² minimum | 200 cm² | ☐ |
| Accuracy | ±1% up to 500 mm/hr | ±2% up to 300 mm/hr | ±5% | ☐ |
| Resolution | 0.2 mm | 0.5 mm | 0.5 mm | ☐ |
| Calibration | Field-adjustable, calibration certificate | Field-adjustable | Factory set | ☐ |
| Output | Reed switch or SDI-12 | Reed switch | Reed switch | ☐ |

  - Water Quality Sensor Specifications (Optional)

| Parameter | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Included? |
| --- | --- | --- | --- | --- |
| Water Temperature | Accuracy: ±0.2°C, Resolution: 0.01°C | Accuracy: ±0.5°C | Accuracy: ±1.0°C | ☐ |
| pH | Accuracy: ±0.1 pH, Range: 2-12, automatic temperature compensation | Accuracy: ±0.2 pH | Accuracy: ±0.5 pH | ☐ |
| Conductivity | Accuracy: ±1% or ±1 µS/cm, Range: 0-200,000 µS/cm, automatic temperature compensation | Accuracy: ±5% | Accuracy: ±10% | ☐ |
| Turbidity | Accuracy: ±2% or 0.5 NTU, Range: 0-1000 NTU, automatic wiping | Accuracy: ±5% | Accuracy: ±10% | ☐ |
| Dissolved Oxygen | Accuracy: ±0.2 mg/L, Range: 0-20 mg/L, automatic calibration | Accuracy: ±0.5 mg/L | Accuracy: ±1.0 mg/L | ☐ |

  - Groundwater Level Specifications

| Parameter | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Included? |
| --- | --- | --- | --- | --- |
| Measurement Range | 0-100m (or site-specific) | 0-50m | 0-30m | ☐ |
| Accuracy | ±0.05% of full scale | ±0.1% of full scale | ±0.25% of full scale | ☐ |
| Resolution | 1 mm | 2 mm | 1 cm | ☐ |
| Sensor Type | Vented pressure transducer preferred | Vented or non-vented with baro | Non-vented with baro | ☐ |
| Cable | Vented, polyurethane jacketed, Kevlar reinforced, marked at 1m intervals | Vented or non-vented, rugged | Standard cable | ☐ |
| Suspension | Stainless steel cable clamp or lock at well cap | Cable clamp | Cable tie | ☐ |

## Spare Parts Requirements by Tier

#### Tier 1 Spare Parts Kit (per 5 stations recommended)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Water level sensor (complete, same as installed) | 2 |  |  |
|  | Rainfall sensor (complete) | 2 |  |  |
|  | Barometric pressure sensor | 2 |  |  |
|  | Water quality sonde (if installed) | 1 |  |  |
| Data Acquisition | Data logger (complete, programmed) | 2 |  |  |
|  | Communication modem (same as installed) | 2 |  |  |
|  | Power supply module | 2 |  |  |
| Power System | Solar panel (same rating) | 2 |  |  |
|  | Charge controller | 2 |  |  |
|  | Battery (deep-cycle, same specification) | 4 |  |  |
| Infrastructure | Sensor cables - full set | 2 sets |  |  |
|  | Connectors and glands kit | 2 sets |  |  |
|  | Desiccant tubes and refill | 10 packs |  |  |
|  | Fuses, surge protectors | 2 sets |  |  |
| Consumables | Desiccant (for vented sensors) | 1 year supply |  |  |
|  | Cable ties (UV-stabilized) | 100 |  |  |
|  | Cleaning kit for water quality sensors | 2 |  |  |
|  | Calibration solutions (if water quality) | 1 set |  |  |
| Tools | Field verification kit (portable level reference, barometer reference) | 1 |  |  |
|  | Laptop with all software | 1 |  |  |
| TOTAL (Tier 1) |  |  |  |  |

#### Tier 2 Spare Parts Kit (per 5 stations recommended)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Water level sensor | 2 |  |  |
|  | Rainfall sensor | 2 |  |  |
|  | Barometric pressure sensor | 1 |  |  |
| Data Acquisition | Data logger | 1 |  |  |
|  | Communication modem | 1 |  |  |
| Power System | Battery | 4 |  |  |
|  | Charge controller | 1 |  |  |
| Infrastructure | Critical cables | 2 sets |  |  |
|  | Connectors and fuses | 2 sets |  |  |
| TOTAL (Tier 2) |  |  |  |  |

#### Tier 3 Spare Parts Kit (per 10 stations recommended)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Water level sensor | 2 |  |  |
|  | Rain gauge (if installed) | 2 |  |  |
| Power | Battery | 4 |  |  |
| Infrastructure | Basic toolkit | 1 |  |  |
| TOTAL (Tier 3) |  |  |  |  |

## Maintenance Schedule

Tier 1 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Monthly | Remote data check, verify transmissions, review data quality, check for sensor drift, plan site visits | 1 hour remote | Hydrologist |
| Quarterly | Site visit: Visual inspection, clean stilling well intakes, check sensor cable, verify staff gauge reading against sensor, clean solar panels, battery check, inspect enclosure seals | 2-3 hours | Technician |
| Bi-annually | Full sensor verification: pressure sensor checks against portable reference, rain gauge calibration check, desiccant replacement, battery terminal cleaning, lightning protection check | 4 hours | Technician |
| Annually | Full system calibration, water quality sensor calibration (if installed), rating curve verification (if changed), firmware updates, thorough cleaning of stilling well | 1 day | Specialist |
| Post-Flood Event | Site inspection after major flood, check for debris damage, sensor positioning, stilling well condition | As needed | Technician |
| 5-Year | Major refurbishment: replace sensor, replace battery, replace critical cables | 1 day | Specialist |

Tier 2 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Monthly (remote) | Remote data check | 30 min | Hydrologist |
| Quarterly | Site visit: Visual inspection, staff gauge check, clean sensor, battery check, solar panel cleaning | 2 hours | Technician |
| Annually | Full site inspection, sensor verification, rain gauge calibration, battery replacement if needed | 4 hours | Technician |

Tier 3 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Bi-annually | Site visit: Visual inspection, staff gauge check (if installed), battery check, clean sensor | 1-2 hours | Technician or Community Observer |
| Annually | Battery replacement, full functional check | 2-3 hours | Technician |

## Field Verification & Calibration

This section defines requirements for two distinct activities. Field Verification is the on-site check of instrument performance against a portable reference. Calibration is a formal, traceable comparison against a national or international standard, typically performed in a laboratory setting. See the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 7.4 for more detail.

  - Field Verification Equipment

The contractor shall provide field verification equipment:

| Item | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Portable Water Level Meter | Accuracy ±2mm, 30m tape, with buzzer/light | Accuracy ±5mm, 20m tape | Not required | ☐ |
| Barometric Reference | Accuracy ±0.2 hPa, calibrated | Accuracy ±0.5 hPa | Not required | ☐ |
| Rain Gauge Calibrator | Graduated cylinder with stand, drip rate control | Graduated cylinder | Not required | ☐ |
| Temperature Reference | Accuracy ±0.2°C, NIST traceable | Accuracy ±0.5°C | Not required | ☐ |
| Conductivity Standard | If water quality installed | If water quality installed | Not required | ☐ |
| pH Buffer Solutions | If water quality installed | If water quality installed | Not required | ☐ |

  - Calibration Requirements

| Component | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Water Level Sensor | Annual factory or laboratory calibration | 2-year calibration | As-failed replacement | ☐ |
| Rain Gauge | Annual field verification, 2-year lab calibration | Annual field verification | As failed | ☐ |
| Water Quality Sensors | Annual factory calibration | 2-year calibration | As failed | ☐ |
| Field Verification Equipment | Annual calibration at accredited lab | Annual calibration | Not applicable | ☐ |

## Warranty & Services

  - Warranty Requirements

| Component | Tier 1 | Tier 2 | Tier 3 | Bidder Offer |
| --- | --- | --- | --- | --- |
| Warranty Period | Minimum 3 years comprehensive | Minimum 2 years comprehensive | Minimum 1 year comprehensive |  |
| Coverage | Parts, labor, shipping both ways | Parts, labor, shipping to vendor | Parts only |  |
| Response Time | <24 hours for critical failure acknowledgment (flood season) | <48 hours for critical failure | <5 working days |  |
| Remote Support | 24/7 hotline during flood season, business hours otherwise | Business hours hotline, remote diagnostics | Email support |  |
| On-site Support | Guaranteed on-site within 7 days for critical failures during flood season | On-site within 14 days if remote cannot resolve | Not included |  |
| Parts Replacement | Advance replacement shipped within 48 hours for critical sensors | Replacement shipped within 1 week | Replacement shipped within 1 month |  |

  - Flood Season Support

For Tier 1 and B stations in flood-prone areas:

| Requirement | Specification | Bidder Offer |
| --- | --- | --- |
| Priority Support | Critical failures prioritized during defined flood season (specify months: ______) |  |
| Rapid Deployment | Temporary replacement sensor available within 48 hours |  |
| Extended Warranty | Flood damage coverage (debris, inundation) optional |  |

  - Service Level Agreement (SLA) Requirements

The contractor shall provide a detailed SLA including:

| SLA Element | Tier 1 Requirement | Tier 2 Requirement | Bidder Offer |
| --- | --- | --- | --- |
| Data availability target | 98% during flood season, 95% annually | 95% annually |  |
| Scheduled maintenance | Maximum 3 days per year, outside flood season, 4 weeks’ notice | Maximum 4 days per year, 2 weeks’ notice |  |
| Critical failure response | <2 hours acknowledgment, <24 hours remote attempt | <24 hours acknowledgment |  |
| Spare parts delivery | <1 week to site | <2 weeks to site |  |

  - Post-Warranty Support Options

The bidder must provide pricing for post-warranty support:

| Option | Description | Annual Cost | Notes |
| --- | --- | --- | --- |
| Option 1: Comprehensive | Full coverage as during warranty |  |  |
| Option 2: Parts only | Spare parts at fixed prices, no labour |  |  |
| Option 3: Time & materials | Support charged as needed |  |  |
| Option 4: Training refresh | Refresher training for new staff |  |  |

## Training Requirements

| Training Type | Audience | Duration | When | Must Include |
| --- | --- | --- | --- | --- |
| Hydrologist training | Water resources staff, forecasters | 2-3 days | Before installation | Data interpretation, flood warning thresholds, rating curves, data quality |
| Technician training | Field technicians | 5 days | During installation | Sensor installation, maintenance, troubleshooting, field verification, safety near water |
| Water Quality training | Lab/water quality staff | 2-3 days | After installation | Sensor calibration, cleaning, data validation (if applicable) |
| Software training | Data managers | 2-3 days | After installation | Database integration, data export, reporting |
| Community observer training | Local observers (Tier 3) | 1 day | At installation | Basic checks, manual readings, reporting issues |

Training materials required:

☐ User manuals (printed + digital) in English

☐ Maintenance manuals with exploded diagrams

☐ Quick reference guides (laminated for field use) – 2 copies per station

☐ Video tutorials (recommended)

☐ Spare parts catalog with part numbers and photos

☐ Troubleshooting flowchart

☐ Preventative maintenance checklist (laminated)

☐ Safety manual (working near water, flood safety)

Training Evaluation:

- Pre- and post-training assessments required

- Practical demonstration of competency (each participant must perform tasks)

- Certificate of completion for each participant

- Training report with recommendations for additional support

## Technical Documentation

The contractor must provide comprehensive technical documentation:

  - System Documentation

| Document | Description | Format | Quantity |
| --- | --- | --- | --- |
| System Design Document | Overall architecture, component specifications, interface definitions, wiring diagrams | Digital (PDF) + 2 hard copies | 2 sets |
| Installation Manual | Step-by-step installation instructions, site preparation requirements, commissioning procedures, alignment specifications | Digital (PDF) + 2 hard copies | 2 sets |
| User Manual | Daily operation procedures, data viewing, basic troubleshooting, maintenance alerts | Digital (PDF) + 2 hard copies | 3 sets |
| Maintenance Manual | Preventive maintenance schedules, troubleshooting guides, repair procedures, part replacement instructions | Digital (PDF) + 2 hard copies | 3 sets |
| Technical Drawings | Mechanical drawings, electrical schematics, cabling diagrams, mast foundation details, sensor mounting details | Digital (PDF + CAD if available) + 2 hard copies | 2 sets |
| Parts Catalog | Complete list of all components with part numbers, suppliers, cross-reference, exploded views | Digital (PDF) + 2 hard copies | 3 sets |
| Software Documentation | User guides, administrator guides, API documentation, data format specifications, configuration guides | Digital (PDF) | 3 sets |
| Rating Curve Guidance | How to develop and update stage-discharge relationships | Digital + 2 hard copies | 2 sets |
| Calibration Procedures | Step-by-step calibration instructions, traceability information, calibration interval recommendations | Digital (PDF) + 2 hard copies | 3 sets |
| Safety Manual | Working near water, flood safety, electrical safety | Digital (PDF) + 2 hard copies | 3 sets |
| Warranty Documentation | Warranty terms, registration, claim procedures, service contact information | Digital (PDF) + 2 hard copies | 3 sets |
| As-Built Documentation | Final as-installed drawings, configuration settings, test results, photographs (within 30 days of completion) | Digital (PDF) + 2 hard copies | 2 sets |

  - Documentation Quality Requirements

- All documentation must be in English

- Language must be clear, avoiding unnecessary jargon

- Diagrams and illustrations must be used extensively

- Troubleshooting guides must include decision trees

- All documents must include revision control and date

- Hard copies must be durable (weather-resistant covers for field manuals)

- Digital copies must be searchable PDFs

- Critical procedures must have bilingual summaries (English + local language) where possible

  - Site-Specific Documentation

For each installed station, the contractor shall provide:

- Station location map and coordinates (WGS84)

- Site photographs (4 directions, close-ups of each component)

- Sensor heights and orientations

- Configuration file for datalogger

- Network and communication settings

- Contact list for support

## Lifecycle Cost Calculation

| Cost Category |  | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | 5-Year Total |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Capital costs | Sensors (level, rain, etc.) |  |  |  |  |  |  |
|  | Datalogger |  |  |  |  |  |  |
|  | Telemetry equipment |  |  |  |  |  |  |
|  | Power system |  |  |  |  |  |  |
|  | Stilling well / infrastructure |  |  |  |  |  |  |
|  | Installation |  |  |  |  |  |  |
|  | Training |  |  |  |  |  |  |
| Recurring costs | Data plans |  |  |  |  |  |  |
|  | Consumables (desiccant, etc.) |  |  |  |  |  |  |
|  | Batteries |  |  |  |  |  |  |
|  | Calibration solutions |  |  |  |  |  |  |
| Maintenance | Spare parts |  |  |  |  |  |  |
|  | Calibration services |  |  |  |  |  |  |
|  | Software support |  |  |  |  |  |  |
|  | Travel for maintenance |  |  |  |  |  |  |
| TOTAL |  |  |  |  |  |  |  |

Notes:

- Include all costs in [USD or local currency]

- Assume normal operation, no catastrophic failures

- Include inflation if multi-year contract

- Separate optional items clearly

## Pacific Requirements

All freshwater & hydrological systems must meet these additional requirements:

| Requirement | Specification | Compliant? |
| --- | --- | --- |
| Corrosion protection | Marine-grade materials for coastal sites, all components corrosion-resistant |  |
| UV resistance | All outdoor components UV-stabilized |  |
| Cyclone resistance | Sensor mounts rated for 200 km/h winds; stilling well secured against debris |  |
| Flood-proofing | Electronics above expected flood level; sensors rated for immersion; cables secured |  |
| Lightning protection | Comprehensive protection; Pacific has high lightning frequency |  |
| Debris tolerance | Sensor selection considers high debris load during floods |  |
| Biofouling resistance | Anti-fouling measures for sensors in warm freshwater (algae growth) |  |
| Rodent protection | Cables protected from rats; enclosures rodent-proof |  |
| Access during floods | Site access considerations for maintenance (may be cut off) |  |
| Remote diagnostics | Critical for sites inaccessible during flood season |  |
| Power autonomy | 30 days minimum for flood-prone sites during wet season |  |

## Compliance Checklist

Before finalising your tender, verify:

| Check | Completed? |
| --- | --- |
| Tier selection matches your operational needs | ☐ |
| All required variables selected | ☐ |
| Rainfall included where needed | ☐ |
| Water quality sensors specified | ☐ |
| System architecture fully defined | ☐ |
| Telemetry appropriate for site (cellular/satellite) | ☐ |
| Power system sized for location (shading, cloudy periods, 30-day autonomy) | ☐ |
| Flood-proofing considered (electronics above flood level) | ☐ |
| Stilling well or mounting specified appropriately | ☐ |
| Staff gauge included for manual verification | ☐ |
| All sensor specifications are included and match the tier selected | ☐ |
| Data transmission requirements specified and appropriate for the site | ☐ |
| Physical infrastructure requirements clear and appropriate for site conditions | ☐ |
| Software and integration requirements specified | ☐ |
| Spare parts kit matches tier requirement and network size | ☐ |
| Maintenance schedule defined and resourced | ☐ |
| Field verification equipment included |  |
| Warranty and SLA requirements specified | ☐ |
| Training requirements fully defined and specified | ☐ |
| Technical documentation requirements included | ☐ |
| Pacific requirements included | ☐ |
| Lifecycle cost worksheet completed | ☐ |
| Budget confirmed for 5-year operation | ☐ |
| Site surveys completed for all locations | ☐ |

## Additional Information for Bidders

  - Site Information (to be completed by NMHS)

| Site | Latitude | Longitude | Elevation | Access | Power | Comms | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |

  - Existing Equipment at Sites

| Site | Existing Logger | Existing Sensors | Compatible? |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |

  - Special Local Requirements

Import duties: ☐ NMHS responsible ☐ Bidder responsible

Customs clearance: ☐ NMHS assists ☐ Bidder responsible

Local taxes: ☐ Applicable ☐ Exempt (specify)

Work permits: ☐ Required ☐ Not required

Site access restrictions: _________________________________