# Procurement ToolKit – Annex B

Part of Weather Ready Pacific Standards and Guidelines Document

Module: Land

Version: 1.0 | March 2026

When to Use This Module

Use this template when procuring:

- Soil moisture and / or soil temperature

- Full agricultural weather station (includes AWS variables)

- Evapotranspiration (ET) station

- Phenology / crop monitoring station

- Landslide early warning station

IMPORTANT NOTICE

This Procurement Toolkit is a guidance document only, developed under the Weather Ready Pacific (WRP) Programme to support standardised procurement of Earth observation equipment across Pacific Island Countries and Territories.

It is not a complete tender document. Users may need to adapt all content to:

- National procurement laws and policies

- Local environmental and workplace safety regulations

- Project-specific requirements and site conditions

- Applicable customs, import, and tax requirements

Professional judgment is required. Users should seek appropriate technical and legal advice before finalising any tender.

Limitation of Liability: The WRP Programme, its implementing partners, and supporting organisations accept no liability for any decisions made or actions taken based on this guidance.

Living Document: This toolkit will be updated periodically. Verify you have the latest version through SPREP or the WRP Programme.

Contact: Weather Ready Pacific Programme, SPREP, Apia, Samoa [include the email address of the WRP programme?]

## System Definition

Project Name: _________________________________

Number of station(s): _____

Location(s): _________________________________

Station types:

☐Soil moisture only _____ stations

☐ Soil moisture & temperature _____ stations

☐ Full agricultural weather station (includes AWS variables) _____ stations

☐ Evapotranspiration (ET) station _____ stations

☐ Phenology / crop monitoring station _____ stations

☐ Landslide early warning station _____ stations

Land use type: ☐ Agriculture (crops) ☐ Pasture / grazing ☐ Forest ☐ Urban ☐ Bare soil

☐ Research ☐ Unknown

Crop type (if applicable): __________________________

Primary purpose: ☐ Drought monitoring ☐ Agricultural advisory

☐ Water resource management ☐ Landslide warning ☐ Climate research

☐ Fire danger warning ☐ Unknown

Power available at site: ☐ Grid ☐ Solar only ☐ Generator ☐ Battery only ☐ Unknown

Cellular coverage: ☐ Yes ☐ No ☐ Partial ☐ Unknown

WIGOS Station Type: ☐ GBON ☐ RBON ☐ National Only

Soil type: ☐ Clay ☐ Loam ☐ Sandy ☐ Volcanic ☐ Peat ☐ Mixed / Unknown

Slope: ☐ Flat (0-5°) ☐ Moderate (5-15°) ☐ Steep (>15°) – for landslide risk ☐ Unknown

Vegetation cover: ☐ Bare soil ☐ Low crops ☐ Tall crops ☐ Grass ☐ Forest

☐ Mixed / Unknown

Livestock present: ☐ Clay ☐ Loam ☐ Sandy ☐ Volcanic ☐ Peat ☐ Mixed / Unknown

## Tier Selection

| Tier | Typical Use Case | NEMS / International Standards | Recommended For |
| --- | --- | --- | --- |
| 1 (Reference) | Climate research, long-term trends, calibration of satellite products, benchmark sites | GCOS, NEMS QC600 | Research stations, national reference sites, climate monitoring network, satellite validation |
| 2 (Synoptic) | Operational drought monitoring, agricultural advisories, water management | NEMS QC500 | Agricultural extension, catchment management, fire danger rating, landslide warning |
| 3 (Basic) | Community-based monitoring, farmer networks, education | NEMS QC400 | Community gardens, schools, farmer cooperatives, citizen science |

Selected Tier: ☐ 1 (Reference) ☐ 2 (Synoptic) ☐ 3 (Basic) ☐ Mixed (specify: _________)

Note: If a single site will contain equipment from different tiers (e.g., a Tier 1 wind sensor with a Tier 3 rain gauge), this must be clearly specified. The procurement documents should list the required tier for each sensor or system component.

** Your selected tier will dictate the rest of the tender

## System Architecture

This section defines how the system is built, how components interact, and how data flows from sensor to user.

  - Data Acquisition & Processing (Datalogger)

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Processor | 32-bit, industrial grade, low power consumption, -40°C to +70°C operating range | 16-bit or 32-bit, industrial grade, -30°C to +60°C | 16-bit, commercial grade, 0°C to +50°C |  |
| Memory | ≥64 MB, ≥1 year data storage at 15-minute intervals (higher frequency during events) | ≥32 MB, ≥6 months data storage | ≥8 MB, ≥3 months data storage |  |
| Analog Channels | ≥8 differential or 16 single-ended (for multiple soil sensors at various depths) | ≥4 differential or 8 single-ended | ≥2 differential or 4 single-ended |  |
| Digital Channels | ≥4 digital I/O, ≥2 serial/RS232, SDI-12 (≥8 sensors) – soil sensor networks require many SDI-12 connections | ≥2 digital I/O, ≥1 serial, SDI-12 (≥4 sensors) | SDI-12 capable |  |
| Pulse Counters | ≥2 pulse counters (for rain gauge, if included) | ≥2 pulse counters | ≥1 pulse counter |  |
| Scan Rate | ≤10 seconds for all channels | ≤30 seconds | ≤60 seconds |  |
| A/D Resolution | 24-bit (for accurate soil moisture measurements) | 16-bit | 16-bit |  |
| Accuracy | ±0.02% of reading | ±0.1% of reading | ±0.25% of reading |  |
| Time Synchronization | GPS-integrated or NTP, automatic daily sync, <1 second accuracy | GPS or NTP, <5 seconds accuracy | Manual or NTP when available |  |
| Real-Time Clock | Battery-backed, <2 minutes/year drift | Battery-backed, <5 minutes/year drift | Battery-backed, <10 minutes/year drift |  |
| Communications Ports | Ethernet, USB, RS-232, RS-485, maintenance port | USB, RS-232, maintenance port | USB or serial, maintenance port |  |
| Data Formats | Native: binary; Export: CSV, JSON, WaterML2 | Export: CSV, JSON | Export: CSV |  |
| Remote Programming | Full remote configuration and firmware update | Remote configuration | Manual only |  |
| Diagnostics | Comprehensive self-diagnostics, watchdog timer, voltage monitoring, sensor health checks | Basic self-diagnostics, watchdog timer | Watchdog timer |  |

Additional Requirements for Data Acquisition for all tiers:

- Soil Sensor Compatibility: Must support multiple soil sensor types (capacitance, TDR, heat pulse, etc.) on same SDI-12 bus

- Multi-depth Capability: Ability to log soil moisture and temperature at multiple depths (typically 5, 10, 20, 50, 100 cm)

- Low Power Design: Critical for remote agricultural sites with small solar panels

- Maintenance Port: Local laptop connection for field interventions

- Enclosure: Weatherproof (IP66), UV-stabilized, with lock

- Rodent/Ant Protection: Sealed cable entries, pest-resistant design

  - Sensor Complement

| Sensor Type | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Soil Moisture | ✓ Mandatory (multiple depths) | ✓ Mandatory (multiple depths) | ✓ Mandatory (single depth) | ☐ |
| Soil Temperature | ✓ Mandatory (multiple depths) | ✓ Mandatory (multiple depths) | ☐ Optional | ☐ |
| Air Temperature | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Relative Humidity | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Rainfall | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Solar Radiation | ✓ Mandatory (for ET) | ✓ Mandatory (for ET) | ☐ Optional | ☐ |
| Wind Speed | ✓ Mandatory (for ET) | ✓ Mandatory (for ET) | ☐ Optional | ☐ |
| Wind Direction | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Atmospheric Pressure | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Leaf Wetness | ☐ Optional (for disease prediction) | ☐ Optional | ☐ Optional | ☐ |
| Evapotranspiration (ET) | ✓ Calculated from above | ✓ Calculated from above | ☐ Estimated | ☐ |
| Phenology Camera | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Tensiometers (soil water potential) | ☐ Optional (research) | ☐ Optional | ☐ Optional | ☐ |

Selected Sensors for This Procurement:

List the specific sensors required:

| Sensor | Quantity | Manufacturer Preference (if any) | Notes |
| --- | --- | --- | --- |

  - Soil Moisture Sensor Type Selection:

| Sensor Type | Advantages | Disadvantages | Suitable for Tier | Selected? |
| --- | --- | --- | --- | --- |
| Capacitance/FDR | Low cost, low power, good accuracy, multiple depths | Soil-specific calibration required, affected by salinity | 1, 2, 3 | ☐ |
| Time Domain Reflectometry (TDR) | High accuracy, less soil-specific, measures dielectric permittivity | Higher cost, higher power, complex electronics | 1, 2 | ☐ |
| Time Domain Transmissometry (TDT) | Good accuracy, lower power than TDR | Higher cost | 1, 2 | ☐ |
| Neutron Probe | Very accurate, large volume measurement | Regulatory issues (radioactive source), not for automation | 1 only (manual) | ☐ |
| Gypsum Blocks | Very low cost, simple | Slow response, limited range, dissolves over time | 3 | ☐ |
| Tensiometers | Measures water potential (plant stress) directly | High maintenance, limited range | 1, 2 (research) | ☐ |

  - Data Transmission & Telemetry

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Communication | Dedicated: ☐ Cellular (4G/LTE) ☐ Satellite (BGAN/Starlink/VSAT) ☐ Radio | ☐ Cellular ☐ Satellite ☐ Radio (choose one) | ☐ Cellular ☐ Satellite ☐ Manual download |  |
| Secondary Communication | ☐ Cellular backup (if primary satellite) ☐ Satellite backup (if primary cellular) ☐ Not required | ☐ Recommend where available | Not required |  |
| Data Protocol | TCP/IP, HTTP/HTTPS, MQTT, FTP/SFTP, push to NMHS server | TCP/IP, HTTP/HTTPS, FTP, push to NMHS server | FTP or email, or manual download |  |
| Transmission Frequency | Real-time: ≤15-minute intervals; Daily summary with hourly data | ≤30-minute intervals | Daily summary or on request |  |
| Alert Transmission | Immediate transmission when thresholds exceeded (e.g., soil moisture below critical level) | Threshold alerts where configured | Manual alert only |  |
| Data Formats | CSV, JSON, WaterML2; Compatible with agricultural databases | CSV, JSON | CSV only |  |
| Data Plan | Minimum 3-year prepaid service included (for satellite) | Minimum 2-year prepaid service included | NMHS provides SIM or manual |  |
| Redundancy | Automatic failover, store-and-forward during outages | Store-and-forward | Store-and-forward only |  |
| Remote Monitoring | Full link quality monitoring, signal strength, data usage, alarms | Basic link monitoring | Not required |  |

Agricultural Alert Integration:

System must be capable of sending alerts for:

- Drought conditions (soil moisture below threshold)

- Frost risk (if applicable)

- Heavy rainfall/erosion risk

- Disease risk (leaf wetness duration)

Data Flow Diagram Requirement regardless of tier:

The bidder must provide a detailed data flow diagram showing:

- Path from sensor → datalogger → transmission → NMHS server / agriculture server

- All data formats at each stage

- Backup paths and failover mechanisms

- Store-and-forward operation during outages

- Points of manual intervention

Integration Testing:

The contractor shall demonstrate successful data transmission to:

☐ NMHS central server / agriculture server

☐ WIS 2.0 node or GTS connection

☐ Forecaster workstation display

  - Power System

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Power Source | ☐ Grid with UPS ☐ Solar (specify below) | ☐ Grid with battery ☐ Solar | ☐ Solar ☐ Battery only |  |
| Solar Panel | Minimum 100W, monocrystalline, marine-grade anodized frame, 25-year lifespan | Minimum 80W, monocrystalline, corrosion-resistant frame | Minimum 50W, polycrystalline |  |
| Charge Controller | MPPT, 20A minimum, temperature compensation, remote monitoring capable | MPPT or PWM, 15A, temperature compensation | PWM, 10A |  |
| Battery Bank | Deep-cycle AGM or Lithium, 100Ah minimum, 5-year design life, vented enclosure | Deep-cycle AGM, 80Ah minimum, 3-year design life | Deep-cycle lead-acid, 50Ah minimum |  |
| Autonomy | Minimum 21 days without sun, full operation | Minimum 14 days without sun, full operation | Minimum 7 days without sun, critical sensors only |  |
| Power Distribution | Dedicated distribution panel, circuit breakers, surge protection on all lines | Fused distribution, surge protection on main line | Fused distribution |  |
| Voltage Regulation | Regulated outputs: 12V, 5V; all sensors protected | Regulated 12V output | Battery voltage direct |  |
| Power Monitoring | Remote monitoring: voltage, current, battery health, charge status, alarms | Local display of voltage and charge status | Visual battery check |  |
| Enclosure | Lockable, weatherproof (IP66), ventilated, corrosion-resistant | Lockable, weatherproof (IP65), ventilated | Weatherproof, basic lock |  |

Important Note for Land Sites:

Agricultural sites often have:

- Shading from crops/trees – may require larger panels or careful siting

- Livestock – enclosures must be animal-proof (fencing required)

- Remote access – longer autonomy needed if sites are infrequently visited

- Dust – solar panels require regular cleaning in dry conditions

Solar Power Sizing Calculation:

The bidder must provide detailed power budget showing:

- System total power consumption (watts)

- Daily energy requirement (watt-hours)

- Solar panel sizing calculation

- Battery bank sizing calculation

- Expected autonomy under worst-case conditions

  - Physical Infrastructure

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Sensor Installation - Soil | Buried sensors at standard depths: 5, 10, 20, 50, 100 cm (or site-specific), undisturbed soil, backfilled carefully, cables protected | Depths: 10, 20, 50 cm, careful installation | Single depth (10-20 cm), simple installation |  |
| Sensor Access | Access tubes for calibration verification, permanent markers for each depth | Access tubes recommended | Surface markers |  |
| Mast / Pole | 3-4m galvanized steel for meteorological sensors, concrete foundation | 2-3m galvanized steel, concrete or ground screw | 2m pole, ground screw |  |
| Instrument Enclosure | Lockable fiberglass or stainless steel, IP66, UV-stabilized, sun shield, passive ventilation, pest-proof screening | Lockable, IP65, UV-stabilized, pest-proof | Lockable, weatherproof |  |
| Fencing / Security | 5m x 5m fenced compound (if livestock present), 2m high chain-link, locked gate, wildlife-proof | Fenced area if livestock present | Locked enclosure only |  |
| Lightning Protection | Full system: lightning rod on mast, down conductor, ground rods (<10 ohm), surge arrestors on all lines | Lightning rod, ground rod, surge arrestors on power | Basic lightning rod or disconnect during storms |  |
| Grounding | <10-ohm ground resistance | <25-ohm resistance | Basic ground rod |  |
| Cabling | UV-resistant, rodent-proof, buried at minimum 30cm depth to avoid cultivation damage, labelled both ends, waterproof glands | UV-resistant, buried, labelled | Outdoor cable, secured |  |
| Cable Entry | Watertight glands, drip loops, pest-proof seals | Watertight glands | Basic cable entry |  |
| Site Markers | Permanent marker with station name, contact details, warning signs | Station marker | Basic marker |  |

Soil Sensor Installation Requirements:

| Depth | Installation Method | Verification |
| --- | --- | --- |
| 5 cm | Horizontal insertion from pit face or vertical from surface | Measure exact depth after installation |
| 10 cm | Horizontal insertion from pit face | Photograph before backfilling |
| 20 cm | Horizontal insertion from pit face | Record GPS location of each sensor |
| 50 cm | Horizontal insertion from pit face or augured hole | Soil description at each depth |
| 100 cm | Augured hole with careful backfill | Bulk density sample recommended |

Important: Soil sensors must be installed in undisturbed soil – digging a large pit and backfilling changes soil properties. Installation from a pit face or using a soil corer with minimal disturbance is required.

Site Preparation Requirements:

The contractor shall provide detailed site preparation guide including:

- Soil pit excavation and backfill procedures (if required)

- Sensor installation protocol for each depth

- Cable trenching depths (minimum 30cm to avoid cultivation)

- Foundation specifications for meteorological mast

- Grounding system installation

- Site access requirements

- Fencing specifications (if livestock present)

- Local regulatory approvals (if any)

  - Software & Data Management

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Logger Configuration Software | Full-featured, graphical interface, offline programming, data retrieval, real-time monitoring | Full-featured, offline programming | Basic configuration tool |  |
| Data Viewing Software | Real-time and historical data display, soil moisture profiles, time series, ET calculation | Real-time display, basic graphing | Not required |  |
| Evapotranspiration (ET) Calculation | Automated ET0 calculation using Penman-Monteith or equivalent, configurable crop coefficients | ET0 calculation | Not required |  |
| Drought Index Software | Soil moisture deficit calculation, comparison with historical records, drought classification | Basic deficit calculation | Not required |  |
| Alert Software | Configurable thresholds for soil moisture, temperature, rainfall; SMS/email alerts | Basic threshold alerts | Not required |  |
| Data Management System | Compatible with agricultural databases, CliDE, automatic data push | Compatible with national database, data push | CSV export for manual import |  |
| WaterML2 Compliance | Full WaterML2 output capability | Preferred | Not required |  |
| Mobile Application | Optional: mobile app for farmers to view soil moisture | Optional | Not required |  |
| Web Interface | Optional: web-based data viewing for farmers and advisors | Optional | Not required |  |
| License Model | Perpetual license, no recurring fees | Perpetual license | Included |  |
| Software Updates | 5 years free updates included | 3 years updates included | As available |  |

  - Integration with Existing Systems

The proposed system must integrate with the following systems:

| System | Integration Requirement | Compliant? |
| --- | --- | --- |
| [e.g., Agricultural Advisory System] | Real-time soil moisture data for irrigation scheduling |  |
| [e.g., Drought Early Warning System] | Soil moisture deficit data, comparison with thresholds |  |
| [e.g., CliDE] | Compatible data formats, automatic transfer |  |
| [e.g., National Hydrology Database] | Soil moisture for runoff modelling |  |
| [e.g., Fire Danger Rating System] | Soil moisture, fuel moisture estimates |  |
| [e.g., Farmer SMS Service] | Automated alerts for irrigation needs |  |

Integration Testing:

The contractor shall demonstrate successful integration during:

- Factory Acceptance Test (simulated environment)

- Site Acceptance Test (live at installation)

- Alert System Test (simulated threshold events)

## Technical Specification by Tier

The specifications in this table are derived from the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 5.2.5 (Tier Sensor Specification by Variable – Land). Refer to the Guide for detailed sensor accuracy and uncertainty data.

  - Soil Moisture Sensor Specifications

Tier 1 - Reference Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Measurement Range | 0-100% Volumetric Water Content (VWC) |  |  |
| Accuracy | ±2% VWC after soil-specific calibration, ±3% with factory calibration |  |  |
| Resolution | 0.1% VWC |  |  |
| Measurement Volume | ≥500 cm³ (representative sample) |  |  |
| Sensor Output | VWC, temperature, bulk electrical conductivity (optional) |  |  |
| Operating Temperature | -20°C to +60°C |  |  |
| Stability | <1% drift per year |  |  |

Special Requirements for Tier 1:

- Soil-specific calibration recommended (samples analysed in laboratory)

- Installation with minimal soil disturbance

- Access tubes for periodic calibration verification

- Co-located with soil temperature sensors at same depths

Tier 2 - Synoptic Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Measurement Range | 0-100% VWC |  |  |
| Accuracy | ±4% VWC with factory calibration |  |  |
| Resolution | 0.1% VWC |  |  |
| Operating Temperature | -10°C to +50°C |  |  |
| Output | SDI-12 or analog |  |  |
| Depth Capability | Multiple depths |  |  |

Tier 3 - Basic Station

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Measurement Range | 0-100% VWC |  |  |
| Accuracy | ±5-8% VWC |  |  |
| Resolution | 0.5-1% VWC |  |  |
| Output | SDI-12 or analog |  |  |
| Depth | Single depth (10-20 cm) |  |  |

  - Soil Temperature Sensor Specifications

| Parameter | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- | --- | --- |
| Measurement Range | -20°C to +60°C | -20°C to +60°C | -10°C to +50°C |  |  |
| Accuracy | ±0.2°C | ±0.3°C | ±0.5°C |  |  |
| Resolution | 0.1°C | 0.1°C | 0.1°C |  |  |
| Sensor Type | Thermistor or RTD | Thermistor | Thermistor |  |  |
| Response Time | <1 minute | <2 minutes | <5 minutes |  |  |
| Output | SDI-12 or 4-20mA | SDI-12 or analog | Analog |  |  |

  - Evapotranspiration (ET) Calculations Requirements

| Parameter | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Compliant? |
| --- | --- | --- | --- | --- |
| Sensor Type | Grid impedance or capacitive | Grid impedance |  | ☐ |
| Measurement | Wet/dry status, duration of wetness | Wet/dry status |  | ☐ |
| Accuracy | ±15 minutes wetness duration | ±30 minutes |  | ☐ |
| Mounting | 45° angle, representative of crop canopy | 45° angle |  | ☐ |

## Spare Parts Requirements by Tier

#### Tier 1 Spare Parts Kit (per 5 stations)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Soil moisture sensor (complete) - each depth type | 2 per depth |  |  |
|  | Soil temperature sensor | 2 per depth |  |  |
|  | Air temperature/RH sensor | 2 |  |  |
|  | Rainfall sensor (complete) | 2 |  |  |
|  | Solar radiation sensor | 1 |  |  |
|  | Wind speed sensor | 1 |  |  |
| Data Acquisition | Data logger (complete, programmed) | 2 |  |  |
|  | Communication modem (same as installed) | 2 |  |  |
| Power System | Solar panel (same rating) | 2 |  |  |
|  | Charge controller | 2 |  |  |
|  | Battery (deep-cycle, same specification) | 4 |  |  |
| Infrastructure | Sensor cables - full set (with waterproof connectors) | 2 sets |  |  |
|  | Connectors and glands kit | 2 sets |  |  |
|  | Fuses, surge protectors | 2 sets |  |  |
| Consumables | Desiccant packs | 20 |  |  |
|  | Cable ties (UV-stabilized) | 200 |  |  |
|  | Calibration solutions (if applicable) | As required |  |  |
| Tools | Field verification kit (soil moisture reference, temperature reference) | 1 |  |  |
|  | Soil auger for re-installation | 1 |  |  |
|  | Installation tool kit | 1 |  |  |
| TOTAL (Tier 1) |  |  |  |  |

#### Tier 2 Spare Parts Kit (per 5 stations)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Soil moisture sensor | 2 per depth |  |  |
|  | Soil temperature sensor | 2 per depth |  |  |
|  | Air temperature/RH sensor | 2 |  |  |
|  | Rainfall sensor | 2 |  |  |
| Data Acquisition | Data logger | 2 |  |  |
|  | Communication modem | 2 |  |  |
| Power System | Battery | 4 |  |  |
|  | Charge controller | 2 |  |  |
| Infrastructure | Critical cables | 2 sets |  |  |
| TOTAL (Tier 2) |  |  |  |  |

#### Tier 3 Spare Parts Kit (per 10 stations)

| Component Category | Item | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- | --- |
| Sensors | Soil moisture sensor | 4 |  |  |
|  | Air temperature sensor (if installed) | 2 |  |  |
|  | Rain gauge (if installed) | 2 |  |  |
| Power | Battery | 8 |  |  |
| Infrastructure | Basic toolkit | 2 |  |  |
| TOTAL |  |  |  |  |

## Maintenance Schedule

Tier 1 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Weekly (remote) | Remote data check, verify transmissions, review soil moisture trends, check battery voltage, check for sensor drift | 30 min | Agrometeorologist |
| Monthly | Site visit: Visual inspection of all equipment, clean solar panels, check enclosure seals, verify against any manual observations | 1-2 hours | Technician |
| Quarterly | Vegetation management (ensure sensors not shaded by crop growth), check cable conditions, battery terminal cleaning, verify sensor depths (soil settling) | 2-3 hours | Technician |
| Bi-annually | Full sensor verification: compare soil moisture readings with gravimetric samples (laboratory analysis), recalibrate if needed, check all connections | 4 hours | Specialist |
| Annually | Full system calibration check, battery capacity test, software updates, sensor replacement if drift detected, soil bulk density sampling | 1 day | Specialist |
| Post-Harvest / Cultivation | Site inspection after cultivation, check for sensor damage, cable exposure, reinstall if necessary | As needed | Technician |
| 5-Year | Major refurbishment: replace all soil sensors (they degrade), replace batteries, replace critical cables, full system recertification | 1 week | Specialist |

Tier 2 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Monthly (remote) | Remote data check | 30 min | Technician |
| Quarterly | Site visit: Visual inspection, vegetation management, clean solar panels, battery check | 1-2 hours | Technician |
| Annually | Full site inspection, sensor verification (field check), battery replacement if needed, calibration check | 4 hours | Technician |

Tier 3 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Quarterly (remote) | Remote data check | 30 min | Technician or Observer |
| Bi-annually | Site visit: Visual inspection, clean sensor, battery check, vegetation management | 1 hour | Technician or Observer |
| Annually | Battery replacement, full functional check | 2 hours | Technician |

## Field Verification & Calibration

This section defines requirements for two distinct activities. Field Verification is the on-site check of instrument performance against a portable reference. Calibration is a formal, traceable comparison against a national or international standard, typically performed in a laboratory setting. See the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 7.4 for more detail.

  - Field Verification Equipment

| Item | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Gravimetric Sampling Kit | Soil auger, sample tins, drying oven (lab), balance (lab) – for calibration verification | Soil auger, sample tins (send to lab) | Not required | ☐ |
| Portable Soil Moisture Meter | Hand-held probe for spot checks, same technology as installed sensors | Hand-held probe | Not required | ☐ |
| Temperature Reference | Accuracy ±0.2°C, NIST traceable | Accuracy ±0.5°C | Not required | ☐ |
| Infiltrometer (optional) | For hydraulic conductivity measurements | Not required | Not required | ☐ |
| Bulk Density Kit | Core samplers, oven, balance | Not required | Not required | ☐ |

  - Calibration Requirements

| Component | Tier 1 | Tier 2 | Tier 3 |
| --- | --- | --- | --- |
| Soil Moisture Sensor | Soil-specific calibration recommended initially; verify every 2 years with gravimetric samples | Factory calibration; spot-check annually | Factory calibration; as failed |
| Soil Temperature Sensor | Annual verification against reference | 2-year verification | As failed |
| Meteorological Sensors | Annual laboratory calibration (as per AWS module) | 2-year calibration | As failed |
| Field Verification Equipment | Annual calibration at accredited lab | Annual calibration | N/A |

Important Note on Soil Moisture Calibration:

Soil moisture sensors measure dielectric permittivity, which is converted to VWC using a calibration equation. Factory calibrations are generic and may be inaccurate for specific soil types. For Tier 1 sites, soil-specific calibration is strongly recommended:

- Collect undisturbed soil samples from each depth at installation

- Measure actual VWC gravimetrically (weigh, dry, reweigh)

- Compare with sensor reading

- Develop site-specific calibration equation

- Apply in datalogger or post-processing

## Warranty & Services

  - Warranty Requirements

| Component | Tier 1 | Tier 2 | Tier 3 | Bidder Offer |
| --- | --- | --- | --- | --- |
| Warranty Period | Minimum 3 years comprehensive | Minimum 2 years comprehensive | Minimum 1 year comprehensive |  |
| Coverage | Parts, labor, shipping both ways | Parts, labor, shipping to vendor | Parts only, return to vendor |  |
| Soil Sensor Specific | Coverage for sensor drift, seal failure, cable damage | Basic coverage | Limited |  |
| Response Time | <24 hours for critical failure acknowledgment | <48 hours for critical failure | <5 working days |  |
| Remote Support | 24/7 hotline, remote diagnostics included | Business hours hotline, remote diagnostics | Email support |  |
| On-site Support | Guaranteed within 14 days for critical failures | Within 30 days if remote cannot resolve | Not included |  |
| Parts Replacement | Advance replacement within 1 week | Replacement within 2 weeks | Replacement within 1 month |  |

  - Agricultural Season Support (Tier 1 and 2)

| Requirement | Specification | Bidder Offer |
| --- | --- | --- |
| Priority Support | Critical failures prioritised during growing season |  |
| Rapid Deployment | Temporary replacement sensor available within 1 week |  |
| Extended Warranty | Coverage for cultivation damage (optional) |  |

  - Service Level Agreement (SLA) Requirements

The contractor shall provide a detailed SLA including:

| SLA Element | Tier 1 Requirement | Tier 2 Requirement | Bidder Offer |
| --- | --- | --- | --- |
| Availability target | 98% data availability (excluding scheduled maintenance) during growing season, 95% annually | 95% data availability annually |  |
| Scheduled maintenance | Maximum 4 days per year, outside growing season, advance notice 4 weeks | Maximum 5 days per year, advance notice 2 weeks |  |
| Helpdesk hours | 24/7 during growing season, business hours otherwise | Business hours (8am-5pm local time) with after-hours emergency contact |  |
| Critical failure response | <2 hours acknowledgment, <24 hours remote attempt | <24 hours acknowledgment |  |
| Spare parts delivery | <1 week to site (within region) | <2 weeks to site |  |
| Software updates | All updates and security patches included for warranty period | Critical security patches included |  |
| Performance reporting | Monthly report on system performance, issues, resolutions | Quarterly report |  |

  - Post-Warranty Support Options

The bidder must provide pricing for post-warranty support:

| Option | Description | Annual Cost | Notes |
| --- | --- | --- | --- |
| Option 1: Comprehensive | Full coverage as during warranty |  |  |
| Option 2: Parts only | Spare parts at fixed prices, no labour |  |  |
| Option 3: Time & materials | Support charged as needed |  |  |
| Option 4: Training refresh | Refresher training for new staff |  |  |

## Training Requirements

| Training Type | Audience | Duration | When | Must Include |
| --- | --- | --- | --- | --- |
| Agrometeorologist training | Agricultural advisors, forecasters | 2-3 days | Before installation | Soil moisture interpretation, ET calculation, crop water requirements, drought indices |
| Technician training | Field technicians | 5 days | During installation | Soil sensor installation techniques, pit excavation, cable burial, maintenance, troubleshooting |
| Soil Calibration training | Technical specialists | 2-3 days | After installation | Gravimetric sampling, lab procedures, developing calibration equations |
| Software training | Data managers | 2-3 days | After installation | Database integration, alert configuration, ET calculation setup |
| Farmer/Community training | End users (Tier 3) | 1 day | At installation | Reading data (SMS/web), understanding soil moisture, reporting issues |
| Livestock safety training | All field staff | 0.5 day | Before field work | Working in paddocks, electric fences, bull safety |

Training materials required:

☐ User manuals (printed + digital) in English

☐ Maintenance manuals with exploded diagrams

☐ Quick reference guides (laminated for field use) – 2 copies per station

☐ Video tutorials (recommended)

☐ Spare parts catalog with part numbers and photos

☐ Gravimetric sampling procedures (Tier 1 only)

☐ Troubleshooting flowchart (laminated)

☐ Preventative maintenance checklist (laminated)

☐ Farmer information sheet (local language recommended)

Training Venue:

- First session (operator): At NMHS or online

- Second session (maintenance/installation): At NMHS, including field installation at 1-2 sites

- Software training: At NMHS computer lab

Training Evaluation:

- Pre- and post-training assessments required

- Practical demonstration of competency (each participant must perform tasks)

- Certificate of completion for each participant

- Training report with recommendations for additional support

## Technical Documentation

The contractor must provide comprehensive technical documentation:

  - System Documentation

| Document | Description | Format | Quantity |
| --- | --- | --- | --- |
| System Design Document | Overall architecture, component specifications, interface definitions, wiring diagrams | Digital (PDF) + 2 hard copies | 2 sets |
| Installation Manual | Step-by-step installation instructions, site preparation requirements, commissioning procedures, alignment specifications | Digital (PDF) + 2 hard copies | 2 sets |
| User Manual | Daily operation procedures, data viewing, basic troubleshooting, maintenance alerts | Digital (PDF) + 2 hard copies | 3 sets |
| Maintenance Manual | Preventive maintenance schedules, troubleshooting guides, repair procedures, part replacement instructions | Digital (PDF) + 2 hard copies | 3 sets |
| Technical Drawings | Mechanical drawings, electrical schematics, cabling diagrams, mast foundation details, sensor mounting details | Digital (PDF + CAD if available) + 2 hard copies | 2 sets |
| Parts Catalog | Complete list of all components with part numbers, suppliers, cross-reference, exploded views | Digital (PDF) + 2 hard copies | 3 sets |
| Software Documentation | User guides, administrator guides, API documentation, data format specifications, configuration guides | Digital (PDF) | 3 sets |
| Calibration Procedures | Step-by-step calibration instructions, traceability information, calibration interval recommendations | Digital (PDF) + 2 hard copies | 3 sets |
| Safety Manual | All safety procedures, hazard identification, emergency response, safe work practices | Digital (PDF) + 2 hard copies | 3 sets |
| Warranty Documentation | Warranty terms, registration, claim procedures, service contact information | Digital (PDF) + 2 hard copies | 3 sets |
| As-Built Documentation | Final as-installed drawings, configuration settings, test results, photographs (within 30 days of completion) | Digital (PDF) + 2 hard copies | 2 sets |

  - Documentation Quality Requirements

- All documentation must be in English

- Language must be clear, avoiding unnecessary jargon

- Diagrams and illustrations must be used extensively

- Troubleshooting guides must include decision trees

- All documents must include revision control and date

- Hard copies must be durable (weather-resistant covers for field manuals)

- Digital copies must be searchable PDFs

- Critical procedures must have bilingual summaries (English + local language) where possible

  - Site-Specific Documentation

For each installed station, the contractor shall provide:

- Station location map and coordinates (WGS84)

- Site photographs (4 directions from mast, close-ups of each component)

- Sensor heights and orientations

- Grounding resistance measurement

- Commissioning test results

- Configuration file for datalogger

- Network and communication settings

- Contact list for support

## Lifecycle Cost Calculation

| Cost Category |  | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | 5-Year Total |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Capital costs | Equipment (sensors, logger, etc.) |  |  |  |  |  |  |
|  | Mast and infrastructure |  |  |  |  |  |  |
|  | Power system |  |  |  |  |  |  |
|  | Telemetry equipment |  |  |  |  |  |  |
|  | Installation |  |  |  |  |  |  |
|  | Training |  |  |  |  |  |  |
| Recurring costs | Data plans (cellular/satellite) |  |  |  |  |  |  |
|  | Consumables (desiccant, etc.) |  |  |  |  |  |  |
|  | Batteries (if not included in capital) |  |  |  |  |  |  |
| Maintenance | Spare parts (planned) |  |  |  |  |  |  |
|  | Calibration services |  |  |  |  |  |  |
|  | Software licenses / support |  |  |  |  |  |  |
|  | Support contract (post-warranty) |  |  |  |  |  |  |
|  | Travel for maintenance |  |  |  |  |  |  |
| TOTAL |  |  |  |  |  |  |  |

Note: Soil sensors have a finite lifespan (typically 5-8 years) and may require replacement. Include sensor replacement in Year 5-6 planning.

Further notes:

- Include all costs in [USD or local currency]

- Assume normal operation, no catastrophic failures

- Include inflation if multi-year contract

- Separate optional items clearly

## Pacific Requirements

All land-based monitoring systems must meet these additional requirements:

| Requirement | Specification | Compliant? |
| --- | --- | --- |
| Corrosion protection | Marine-grade materials (316 stainless steel, marine aluminium, or powder-coated galvanized steel), sealed connectors, tropical coating on all outdoor components |  |
| UV resistance | All outdoor components UV-stabilized (cabling, enclosures, plastic parts) - minimum 5-year UV life |  |
| Cyclone resistance | Mast and sensors rated to 250 km/h (Tier 1/B), 180 km/h (Tier 3); all components securely fastened |  |
| Salt spray protection | All outdoor components sealed to IP66 minimum, marine-grade fasteners (stainless steel), no exposed mild steel |  |
| Waterproofing | IP66 for all outdoor enclosures, IP67 for connectors, sealed cable entries with glands and drip loops |  |
| Humidity tolerance | 0-100% condensing environment, all electronics conformally coated or protected |  |
| Temperature range | Operational: 0°C to +50°C (typical), survival: -10°C to +60°C |  |
| Mould/fungus resistance | All materials treated or selected for tropical resistance, no organic materials in enclosures |  |
| Insect/pest protection | Screened ventilation openings, sealed cable entries, ant-proof considerations |  |
| Lightning protection | Complete system protection with redundancy, compliance with IEC 62305 or equivalent, <10-ohm ground |  |
| Power autonomy | 21 days minimum for remote agricultural sites (may be visited infrequently) |  |
| Solar panel orientation | Adjustable mounts for optimal angle at latitude (typically 0-15° in tropics) |  |
| Data retrieval during outage | Store-and-forward capability for minimum period equal to power autonomy |  |
| Local supportability | Common connectors, standard tools for maintenance, no proprietary fasteners |  |
| Fallow period survival | System must survive periods of no site access between growing seasons |  |
| High soil temperatures | Tropical soils can exceed 50°C at surface; sensors must be rated accordingly |  |
| High soil salinity | Coastal agricultural areas may have saline soils; sensors must be salinity-tolerant |  |
| Volcanic soils | Some Pacific soils have unusual electrical properties; sensor selection should consider this |  |

## Compliance Checklist

Before finalising your tender, verify:

| Check | Completed? |
| --- | --- |
| Tier selection matches your operational needs | ☐ |
| All required variables selected | ☐ |
| System architecture fully defined | ☐ |
| All sensor specifications are included and match the tier selected | ☐ |
| Data transmission requirements specified and appropriate for the site | ☐ |
| Power system defined correctly for the location | ☐ |
| Livestock presence considered (fencing requirements) | ☐ |
| Multiple depths specific for Tier 1/B | ☐ |
| ET calculation requirements defined (if applicable) | ☐ |
| Physical infrastructure requirements clear and appropriate for cite conditions | ☐ |
| Software and integration requirements specified | ☐ |
| Spare parts kit matches tier requirement and network size | ☐ |
| Maintenance schedule defined and resourced, including vegetation management | ☐ |
| Field verification equipment included | ☐ |
| Warranty and SLA requirements specified | ☐ |
| Training requirements fully defined and specified (including farmer training if applicable) | ☐ |
| Technical documentation requirements included | ☐ |
| Pacific requirements included | ☐ |
| Lifecycle cost worksheet completed (5-year TCO) | ☐ |
| Budget confirmed for capital AND 5-year operation | ☐ |
| Site surveys completed for all locations | ☐ |

## Additional Information for Bidders

  - Site Information (to be completed by NMHS)

| Site | Latitude | Longitude | Land Use | Soil Type | Slope | Livestock? | Elevation | Access | Power | Comms | Crop Type | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |  |  |  |  |  |

  - Existing Equipment at Sites

| Site | Existing Logger | Existing Sensors | Compatible? |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |

  - Soil Sampling Requirements (Tier 1 only)

For each Tier 1 site, the contractor shall:

- Collect undisturbed soil samples from each sensor depth (minimum 3 replicates)

- Analyse for: bulk density, field capacity, wilting point, saturation percentage, texture, organic matter

- Provide laboratory report with results

- Develop soil-specific calibration equation for each depth

  - Special Local Requirements

Import duties: ☐ NMHS responsible ☐ Bidder responsible

Customs clearance: ☐ NMHS assists ☐ Bidder responsible

Landowner permission required: ☐ Yes ☐ No ☐ Unknown

Local taxes: ☐ Applicable ☐ Exempt (specify)

Work permits: ☐ Required ☐ Not required

Cultural site considerations: ☐ Yes ☐ No ☐ Unknown

Farmer engagement required: ☐ Yes ☐ No ☐ Unknown

Site access restrictions: _________________________________