# Procurement ToolKit – Annex B

Part of Weather Ready Pacific Standard, Optimisation & Maintenance Guidelines Document

Module: Upper Air Observing Systems

Version: 1.0 | March 2026

When to Use This Module

Use this template when procuring:

- Radiosonde ground systems (receivers, processing)

- Radiosondes (consumables)

- Pilot balloon systems

- Wind profilers

- Upper-air processing software

IMPORTANT NOTICE

This Procurement Toolkit is a guidance document only, developed under the Weather Ready Pacific (WRP) Programme to support standardised procurement of Earth observation equipment across Pacific Island Countries and Territories.

It is not a complete tender document. Users may need to adapt all content to:

- National procurement laws and policies

- Local environmental and workplace safety regulations

- Project-specific requirements and site conditions

- Applicable customs, import, and tax requirements

Professional judgment is required. Users should seek appropriate technical and legal advice before finalising any tender.

Limitation of Liability: The WRP Programme, its implementing partners, and supporting organisations accept no liability for any decisions made or actions taken based on this guidance.

Living Document: This toolkit will be updated periodically. Verify you have the latest version through SPREP or the WRP Programme.

Contact: Weather Ready Pacific Programme, SPREP, Apia, Samoa [include the email address of the WRP programme?]

## System Definition

Project Name: _________________________________

Number of system(s): _____ (Ground stations)

Number of soundings per year: _____ (e.g., 730 for 2/day)

Location(s): _________________________________

Site types: ☐ Coastal ☐ Inland ☐ Mountain ☐ Urban ☐ Airport

Power available at site: ☐ Grid ☐ Solar only ☐ No power available ☐ Unknown

Cellular coverage: ☐ Yes ☐ No ☐ Partial ☐ Unknown

Gas availability: ☐ Hydrogen ☐ Helium ☐ Neither (to be supplied)

WIGOS Station Type: ☐ GBON ☐ RBON ☐ National Only

## Tier Selection

| Tier | Typical Use Case | WMO Classification | Recommended For |
| --- | --- | --- | --- |
| 1 (Reference) | Climate monitoring, research, model assimilation | Class 1 | Capital cities, major airports, climate reference sites |
| 2 (Synoptic) | Operational forecasting, aviation | Class 2 | Secondary airports, forecast offices |
| 3 (Basic) | Pilot balloons, wind-only profiling | Class 3-4 | Remote sites, wind shear monitoring |

Selected Tier: ☐ 1 (Reference) ☐ 2 (Synoptic) ☐ 3 (Basic) ☐ Mixed (specify: _________)

Note: If a single site will contain equipment from different tiers (e.g., a Tier 1 wind sensor with a Tier 3 rain gauge), this must be clearly specified. The procurement documents should list the required tier for each sensor or system component.

** Your selected tier will dictate the rest of the tender

## System Architecture

This section defines how the system is built, how components interact, and how data flows from sensor to user.

  - Data Acquisition & Processing

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Ground Receiver | Dual-receiver redundancy, automatic switchover, remote monitoring | Single receiver with hot-standby capability | Single receiver, manual operation |  |
| Data Processor | Dedicated processing computer (industrial grade), real-time processing, automatic data formatting | Standard computer with processing software | Basic processing unit or manual calculation |  |
| Processing Software | Full suite: acquisition, QC, formatting, archiving, dissemination, BUFR generation | Standard package: acquisition, BUFR generation, basic QC | Wind-only calculation software |  |
| Data Storage | Minimum 5 years raw data, 10 years processed data, automatic backup | Minimum 2 years data storage | Minimum 1 year data storage |  |
| Local Interface | Full system control, diagnostics, data viewing, manual override | Basic control and monitoring | Simple readout display |  |
| Remote Access | Full remote diagnostics, configuration, data retrieval | Remote data access only | Not required |  |
| Analog Channels | ≥4 analog inputs for auxiliary sensors | ≥2 analog inputs | Not required |  |
| Digital Channels | ≥4 digital I/O, ≥2 serial/RS232 | ≥2 digital I/O, ≥1 serial | Not required |  |
| Scan Rate | ≥1 Hz (1-second data) | ≥0.5 Hz (2-second data) | ≥0.2 Hz (5-second data) |  |
| A/D Resolution | 24-bit where applicable | 16-bit | Not applicable |  |
| Accuracy | ±0.02% of reading | ±0.1% of reading | Not applicable |  |
| Time Synchronization | GPS-disciplined oscillator, <1ms accuracy | GPS synchronization | Manual or GPS |  |
| Real-Time Clock | Battery-backed, <2 minutes/year drift | Battery-backed, <5 minutes/year drift | Battery-backed |  |
| Communications Ports | Ethernet, USB, RS-232, RS-485, maintenance port | USB, RS-232, maintenance port | USB or serial |  |
| Remote Programming | Full remote configuration and firmware update | Remote configuration | Manual only |  |
| Diagnostics | Comprehensive self-diagnostics, watchdog timer, signal quality monitoring | Basic self-diagnostics | Manual checks |  |

Additional Requirements for Data Acquisition for all tiers:

- Metadata Capture: Automatic capture of launch time, location, sonde serial number, operator ID

- Maintenance Port: Local laptop connection for field interventions

- Data Format: Raw data in open format, processed data in WMO BUFR/CREX

  - Sensor Complement

| Sensor Type | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Radiosonde (GPS) | ✓ Mandatory | ✓ Mandatory | ☐ Optional | ☐ |
| Radiosonde (Standard) | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Pilot Balloon Kit | ☐ Optional | ☐ Optional | ✓ Mandatory | ☐ |
| Wind Profiler | ☐ Optional | ☐ Optional | ☐ Optional | ☐ |
| Ground Check Equipment | ✓ Mandatory | ✓ Mandatory | ✓ Mandatory | ☐ |
| Sounding Processing Software | ✓ Mandatory | ✓ Mandatory | ✓ Mandatory | ☐ |
| Backup Receiver | ✓ Mandatory | ☐ Optional | ☐ Optional | ☐ |

Selected Sensors for This Procurement:

List the specific sensors required:

| Sensor | Quantity | Manufacturer Preference (if any) | Notes |
| --- | --- | --- | --- |

  - Data Transmission & Telemetry

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Communication | Dedicated: ☐ Cellular (4G/LTE) ☐ Satellite (BGAN/Starlink/VSAT) ☐ Broadband ☐ Radio | ☐ Cellular ☐ Satellite ☐ Broadband ☐ Radio (choose one) | ☐ Cellular ☐ Satellite ☐ Manual download |  |
| Secondary Communication | ☐ Cellular backup (if primary satellite) ☐ Satellite backup (if primary cellular) ☐ Iridium SBD | ☐ Cellular where available ☐ Not required | Not required |  |
| Data Protocol | TCP/IP, HTTP/HTTPS, MQTT, FTP/SFTP, push to NMHS server | TCP/IP, HTTP/HTTPS, FTP, push to NMHS server | FTP or email, or manual download |  |
| Transmission Frequency | Real-time: each sounding level as measured; Full sounding immediately after landing | Full sounding within 5 minutes of landing | Daily summary |  |
| Data Formats | BUFR (mandatory for GBON), CSV, JSON; SYNOP generation | CSV, JSON; SYNOP generation preferred | CSV only |  |
| Data Plan | Minimum 3-year prepaid service included (for satellite) | Minimum 2-year prepaid service included | NMHS provides SIM or manual |  |
| Redundancy | Automatic failover between primary/secondary, store-and-forward during outages | Manual switchover if backup exists; store-and-forward | Store-and-forward only |  |
| Remote Monitoring | Full link quality monitoring, signal strength, data usage, alarms | Basic link monitoring | Not required |  |

Data Flow Diagram Requirement regardless of tier:

The bidder must provide a detailed data flow diagram showing:

- Path from sensor → receiver → processor → transmission → NMHS server / CliDE

- All data formats at each stage

- Backup paths and failover mechanisms

- Points of manual intervention

Integration Testing:

The contractor shall demonstrate successful data transmission to:

☐ NMHS central server / CliDE

☐ WIS 2.0 node or GTS connection

☐ Forecaster workstation display

  - Power System

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Primary Power Source | ☐ Grid with UPS ☐ Generator backup (if remote) | ☐ Grid with UPS ☐ Solar (if remote) | ☐ Solar ☐ Battery only |  |
| Grid Power Requirements | Stable 220/240V AC, 50/60Hz, with surge protection | Stable 220/240V AC | Not applicable |  |
| UPS | Online UPS, 8-hour runtime at full load, generator connection point | Line-interactive UPS, 4-hour runtime | Not required |  |
| Generator (if remote) | Automatic standby generator, 5kVA minimum, weekly test capability | Manual start generator if required | Not required |  |
| Solar Option (if off grid) | Minimum 1000W solar array, MPPT controller, 400Ah battery bank, 3-day autonomy | Minimum 600W solar, 200Ah battery, 2-day autonomy | Minimum 200W solar, 100Ah battery |  |
| Power Distribution | Dedicated circuit breaker panel, surge protection on all incoming lines | Surge protection on main line | Basic surge protection |  |
| Power Monitoring | Remote monitoring: voltage, current, generator status, alarms | Local monitoring | Visual check |  |
| Enclosure/Shelter | Climate-controlled equipment shelter, secure, ventilated | Equipment cabinet, ventilated | Weatherproof enclosure |  |

Note for Upper Air Systems: Upper air systems have significantly higher power consumption than AWS due to radar/receiver systems, processing computers, and gas generation equipment. Grid power with UPS is strongly preferred. Solar options are only recommended for remote Tier 3 sites with low launch frequency.

Sizing Calculation Requirement:

The bidder must provide detailed power budget showing:

- System total power consumption (watts) by component

- UPS sizing calculation with runtime

- Generator sizing (if applicable)

- For solar options: panel sizing, battery bank, autonomy calculation

  - Physical Infrastructure

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Equipment Shelter | Dedicated room or walk-in enclosure, climate-controlled, secure | Equipment cabinet in existing building, ventilated | Weatherproof enclosure |  |
| Antenna Mast | 3-5m tower, galvanized steel, guyed, lightning rod, anti-corrosion coating | 2-3m mast, galvanized, guyed | 1-2m portable mast or tripod |  |
| Gas Storage | Separate ventilated enclosure, compliant with safety regulations, gas detection | Secure ventilated area, gas detection | Portable gas cylinders with safety caps |  |
| Lightning Protection | Full system: air terminals, down conductors, grounding, surge arrestors on all lines | Basic lightning rod and grounding | Portable system: disconnect during storms |  |
| Grounding | <10 ohm ground resistance, separate ground for power and signal | <25 ohm ground resistance | Basic grounding |  |
| Cabling | UV-resistant, rodent-proof, armoured where exposed, all cables in conduit | UV-resistant, protected where exposed | Standard outdoor cable |  |
| Fencing / Security | Secure compound fencing, locked gate, security lighting, CCTV recommended | Locked equipment area | Locked enclosure |  |
| Safety Equipment | Gas detectors, PPE, emergency shower (for hydrogen), fire extinguisher, first aid kit | Gas detectors, PPE, fire extinguisher | Basic safety equipment |  |

Site Preparation Requirements:

The contractor shall provide detailed site preparation guide including:

- Foundation specifications for mast/tower

- Shelter dimensions and requirements

- Cable trenching requirements

- Grounding system installation

- Safety clearances

- Local regulatory approvals needed (frequency licenses, gas storage permits)

  - Software & Data Management

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Acquisition Software | Real-time data display, quality checks, alarm system | Real-time display, basic QC | Simple readout |  |
| Processing Software | Automatic TEMP/BUFR generation, data archiving, climatological calculations | TEMP/BUFR generation, archiving | Wind calculation only |  |
| Monitoring Software | Remote system health monitoring, alerts, diagnostics | Remote data availability monitoring | Not required |  |
| Database Integration | Direct integration with CliDE or NMHS database, automatic data push | Compatible with CliDE, manual or automatic push | CSV export for manual import |  |
| WIS 2.0 Compliance | Full WIS 2.0 node capability or direct push to regional node | BUFR generation for WIS 2.0 upload | Not required |  |
| Metadata Management | Automatic WIGOS metadata generation, OSCAR/Surface compatible | Manual metadata entry supported | Basic site metadata |  |
| License Model | Perpetual license, no recurring fees | Perpetual license | Included |  |
| Software Updates | 5 years free updates included | 3 years updates included | As available |  |

  - Integration with Existing Systems

The proposed system must integrate with the following systems:

| System | Integration Requirement | Compliant? |
| --- | --- | --- |
| [e.g., CliDE] | Automatic data push in required format, real-time |  |
| [e.g., WIS 2.0 Node] | BUFR generation, automatic transfer |  |
| [e.g., Forecaster Workstation] | Data available in real-time for display |  |
| [e.g., GTS] | TEMP/BUFR messages to national GTS connection |  |

Integration Testing:

The contractor shall demonstrate successful integration during:

- Factory Acceptance Test (simulated)

- Site Acceptance Test (live)

## Technical Specification by Tier

The specifications in this table are derived from the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 5.2.2 (Tier Sensor Specification by Variable – Upper Air). Refer to the Guide for detailed sensor accuracy and uncertainty data.

Tier 1 - Reference Upper Air System

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Temperature sensor | Accuracy: ±0.2°C, Resolution: 0.1°C, Range: -90°C to +60°C |  |  |
| Pressure sensor | Accuracy: ±0.5 hPa, Resolution: 0.1 hPa, Range: 5-1080 hPa |  |  |
| Humidity sensor | Accuracy: ±3% RH, Resolution: 1%, Range: 0-100% |  |  |
| Wind measurement | Accuracy: ±0.5 m/s, Resolution: 0.1 m/s, GPS-based |  |  |
| Sampling rate | ≥1 Hz (1-second data) |  |  |
| Data transmission | Real-time via GTS/WIS 2.0, BUFR format |  |  |
| Ground system | Full redundancy (receiver, processing, power) |  |  |

Special Requirements for Tier 1:

- Dual receiver capability for backup

- Uninterruptible Power Supply (UPS) with 8-hour backup

- Generator connection point

- Lightning protection to international standards

- Remote monitoring and diagnostics

Tier 2 - Synoptic Upper Air System

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Temperature sensor | Accuracy: ±0.5°C, Resolution: 0.1°C, Range: -90°C to +60°C |  |  |
| Pressure sensor | Accuracy: ±1.0 hPa, Resolution: 0.1 hPa, Range: 5-1080 hPa |  |  |
| Humidity sensor | Accuracy: ±5% RH, Resolution: 1%, Range: 0-100% |  |  |
| Wind measurement | Accuracy: ±1.0 m/s, Resolution: 0.1 m/s, GPS or radar-tracked |  |  |
| Sampling rate | ≥0.5 Hz (2-second data) |  |  |
| Data transmission | Real-time via GTS/WIS 2.0, BUFR format |  |  |
| Ground system | Single receiver with critical spares |  |  |

Special Requirements for Tier 2:

- UPS with 4-hour backup

- Basic lightning protection

- Remote monitoring capability

Tier 3 - Basic Upper Air System (Pilot Balloon / Wind Profiler)

| Parameter | Specification | Compliance (Y/N) | Evidence |
| --- | --- | --- | --- |
| Wind measurement | Accuracy: ±2.0 m/s or ±10°, Resolution: 0.5 m/s, 5° |  |  |
| Tracking method | Optical theodolite, GPS, or radar |  |  |
| Sampling rate | ≥0.2 Hz (5-second data) |  |  |
| Data output | Wind speed, direction, shear |  |  |
| Ground system | Portable or fixed installation |  |  |

Special Requirements for Tier 3:

- Battery-operated option for remote sites

- Simple, rugged design

- Minimal maintenance requirements

Consumables (Radiosondes)

| Requirement | Specification | Quantity | Notes |
| --- | --- | --- | --- |
| Sonde type | ☐ GPS radiosonde ☐ Standard radiosonde ☐ Winds-only |  |  |
| Shelf life | Minimum 24 months from delivery |  |  |
| Storage conditions | 0°C to +60°C, non-condensing |  |  |
| Battery life | Minimum 90 minutes operation at 0°C |  |  |
| Frequency | 400.15 MHz or as licensed locally |  |  |
| Calibration | Each sonde supplied with individual calibration data |  |  |

Initial stock requirement: _____ sondes (minimum 12 months’ supply)

Annual consumption estimate: _____ sondes

Gas Supply (for Balloons)

| Requirement | Specification | Quantity | Notes |
| --- | --- | --- | --- |
| Gas type | ☐ Hydrogen ☐ Helium |  |  |
| Supply method | ☐ Cylinders ☐ Generator ☐ On-site production |  |  |
| Safety equipment | Gas detectors, PPE, training |  |  |
| Storage | Secure, ventilated, compliant with local regulations |  |  |

## Spare Parts Requirements by Tier

#### Tier 1 Spare Parts Kit (per system)

| Component | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- |
| Radiosonde receiver (complete) | 1 unit |  |  |
| Processing computer (with software) | 1 unit |  |  |
| GPS antenna | 1 unit |  |  |
| Cables - full set | 2 sets |  |  |
| Power supply / UPS | 1 unit |  |  |
| Battery bank | 1 set |  |  |
| Lightning protection modules | 2 sets |  |  |
| Antenna mast components | 1 kit |  |  |
| TOTAL (Tier 1) |  |  |  |

#### Tier 2 Spare Parts Kit (per system)

| Component | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- |
| Radiosonde receiver (PCB/critical module) | 1 unit |  |  |
| Processing computer (basic) | 1 unit |  |  |
| GPS antenna | 1 unit |  |  |
| Critical cables | 1 set |  |  |
| Power supply | 1 unit |  |  |
| TOTAL (Tier 2) |  |  |  |

#### Tier 3 Spare Parts Kit (per system)

| Component | Minimum Quantity | Unit Price | Total |
| --- | --- | --- | --- |
| Replacement bulbs/optics (if applicable) | 2 sets |  |  |
| Basic toolkit | 1 set |  |  |
| Replacement bulbs/optics (if applicable) | 2 sets |  |  |
| TOTAL (Tier 3) |  |  |  |

## Maintenance Schedule

Tier 1 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Before each launch | Check gas pressure, test receiver, verify GPS, prepare sonde | 30 min | Technician |
| Weekly | Clean antenna, check connections, test backup power, download logs | 2 hours | Technician |
| Monthly | Full system calibration check, antenna alignment, generator test | 4 hours | Specialist |
| Quarterly | Gas system inspection, safety equipment check, spare parts inventory | 1 day | Specialist |
| Annually | Full system calibration, preventive maintenance, software updates | 2-3 days | Manufacturer |
| 5-Year | Major refurbishment: replace critical components, full system recertification | 1 week | Manufacturer |

Tier 2 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Before each launch | Check gas pressure, test receiver | 20 min | Technician |
| Monthly | Clean antenna, check connections, test backup power | 2 hours | Technician |
| Quarterly | System check, battery maintenance | 4 hours | Technician |
| Annually | Calibration check, preventive maintenance | 1 day | Specialist |

Tier 3 Maintenance

| Frequency | Tasks | Estimated Time | By Whom |
| --- | --- | --- | --- |
| Before each use | Visual check, battery test | 15 min | Technician |
| Quarterly | Clean optics, check alignment | 1 hour | Technician |
| Annually | Full service | 4 hours | Specialist |

## Field Verification & Calibration

This section defines requirements for two distinct activities. Field Verification is the on-site check of instrument performance against a portable reference. Calibration is a formal, traceable comparison against a national or international standard, typically performed in a laboratory setting. See the WRP Guidance Document: Standardisation, Optimisation and Maintenance of Earth Observation Equipment in the Pacific Islands, Section 7.4 for more detail.

  - Field Verification Equipment

The contractor shall provide field verification equipment:

| Item | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Test Radiosonde | Known good reference sonde | Known good reference sonde | Not required | ☐ |
| Signal Generator | For receiver testing | For receiver testing | Not required | ☐ |
| GPS Test Unit | For antenna verification | For antenna verification | Not required | ☐ |
| Gas Leak Detector | Mandatory for safety | Mandatory for safety | Not required | ☐ |
| Calibration Standards | Temperature/pressure reference | Temperature reference | Not required | ☐ |

  - Calibration Requirements

| Component | Tier 1 | Tier 2 | Tier 3 | Included? |
| --- | --- | --- | --- | --- |
| Ground Receiver | Annual factory calibration | 2-year calibration | As failed | ☐ |
| Test Equipment | Annual accredited calibration | Annual calibration | Not applicable | ☐ |
| Gas System | Annual certification | Annual certification | As required | ☐ |

## Warranty & Services

  - Warranty Requirements

| Component | Tier 1 Specification | Tier 2 Specification | Tier 3 Specification | Bidder Response |
| --- | --- | --- | --- | --- |
| Warranty Period | Minimum 3 years comprehensive | Minimum 2 years comprehensive | Minimum 1 year comprehensive |  |
| Coverage | Parts, labour, shipping both ways | Parts, labour, shipping to vendor | Parts only |  |
| Response Time | <24 hours for critical failure acknowledgment | <48 hours for critical failure acknowledgment | <5 working days |  |
| Remote Support | 24/7 hotline, remote diagnostics included | Business hours hotline, remote diagnostics | Email support |  |
| On-site Support | Guaranteed on-site within 7 days for critical failures | On-site within 14 days if remote cannot resolve | Not included |  |
| Parts Replacement | Advance replacement shipped within 48 hours | Replacement shipped within 1 week | Replacement shipped within 1 month |  |

  - Service Level Agreement (SLA) Requirements

The contractor shall provide a detailed SLA including:

| SLA Element | Tier 1 Requirement | Tier 2 Requirement | Bidder Offer |
| --- | --- | --- | --- |
| Availability target | 99.5% uptime (excluding scheduled maintenance) | 98% uptime |  |
| Scheduled maintenance | Maximum 2 days per year, advance notice 4 weeks | Maximum 3 days per year, advance notice 2 weeks |  |
| Helpdesk hours | 24/7/365 | Business hours (8am-5pm local time) |  |
| Critical failure response | <2 hours acknowledgment, <24 hours remote attempt | <24 hours acknowledgment |  |
| Spare parts delivery | <1 week to site | <2 weeks to site |  |
| Software updates | All updates and security patches included | Critical security patches included |  |
| Performance reporting | Monthly report on system performance, issues, resolutions | Quarterly report |  |

  - Post-Warranty Support Options

The bidder must provide pricing for post-warranty support:

| Option | Description | Annual Cost | Notes |
| --- | --- | --- | --- |
| Option 1: Comprehensive | Full coverage as during warranty |  |  |
| Option 2: Parts only | Spare parts at fixed prices, no labour |  |  |
| Option 3: Time & materials | Support charged as needed |  |  |
| Option 4: Training refresh | Refresher training for new staff |  |  |

## Training Requirements

| Training Type | Audience | Duration | When | Must Include |
| --- | --- | --- | --- | --- |
| Operator training | Launch technicians | 3-5 days | Before first use | Hands-on launches, troubleshooting, safety |
| Maintenance training | Technical staff | 5 days | During installation | System architecture, diagnostics, repairs |
| Software training | Data processors | 2-3 days | After installation | Data processing, GTS transmission, QC |
| Safety training | All staff | 1 day | Before first use | Gas handling, electrical safety, emergency procedures |
| Train-the-trainer | Senior technicians | 3 days | Optional | How to train new staff |

Training materials required:

☐ User manuals (printed + digital) in English

☐ Maintenance manuals with exploded diagrams

☐ Quick reference guides (laminated for field use)

☐ Video tutorials (recommended)

☐ Spare parts catalog with part numbers

☐ Troubleshooting flowchart

☐ Safety data sheet for all consumables

Training Venue:

- First session (operator): At NMHS or online

- Second session (maintenance/installation): At NMHS, including field installation at 1-2 sites

- Software training: At NMHS computer lab

Training Evaluation:

- Pre- and post-training assessments required

- Practical demonstration of competency (each participant must perform tasks)

- Certificate of completion for each participant

- Training report with recommendations for additional support

## Technical Documentation

The contractor must provide comprehensive technical documentation:

| Document | Description | Format | Quantity |
| --- | --- | --- | --- |
| System Design Document | Overall architecture, component specifications, interface definitions | Digital + 2 hard copies | 1 set |
| Installation Manual | Step-by-step installation instructions, site preparation requirements, commissioning procedures | Digital + 2 hard copies | 1 set |
| User Manual | Daily operation procedures, launch preparation, data recording | Digital + 2 hard copies | 2 sets |
| Maintenance Manual | Preventive maintenance schedules, troubleshooting guides, repair procedures | Digital + 2 hard copies | 2 sets |
| Technical Drawings | Mechanical drawings, electrical schematics, cabling diagrams, antenna patterns | Digital (CAD if available) + hard copies | 1 set |
| Parts Catalog | Complete list of all components with part numbers, suppliers, cross-reference | Digital + 2 hard copies | 2 sets |
| Software Documentation | User guides, administrator guides, API documentation, data formats | Digital | 2 sets |
| Calibration Procedures | Step-by-step calibration instructions, traceability information | Digital + 2 hard copies | 2 sets |
| Safety Manual | All safety procedures, hazard identification, emergency response | Digital + 2 hard copies | 2 sets |
| Warranty Documentation | Warranty terms, registration, claim procedures | Digital + 2 hard copies | 2 sets |
| As-Built Documentation | Final as-installed drawings, configurations, photos | Digital + 2 hard copies | 2 sets |

  - Documentation Quality Requirements

- All documentation must be in English

- Language must be clear, avoiding unnecessary jargon

- Diagrams and illustrations must be used extensively

- Troubleshooting guides must include decision trees

- All documents must include revision control and date

- Hard copies must be durable (weather-resistant covers for field manuals)

- Digital copies must be searchable PDFs

  - As-Built Documentation

Within 30 days of installation completion, the contractor shall provide:

- "As-built" drawings showing actual installation

- Final configuration settings for all equipment

- Network addresses and passwords

- Test results from commissioning

- Photographs of installed system

- Any deviations from original design

## Lifecycle Cost Calculation

| Cost Category |  | Year 1 | Year 2 | Year 3 | Year 4 | Year 5 | 5-Year Total |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Capital costs | Ground system |  |  |  |  |  |  |
|  | Installation |  |  |  |  |  |  |
|  | Training |  |  |  |  |  |  |
| Recurring costs | Radiosondes |  |  |  |  |  |  |
|  | Balloons |  |  |  |  |  |  |
|  | Gas (H₂/He) |  |  |  |  |  |  |
|  | Batteries |  |  |  |  |  |  |
| Maintenance | Spare parts |  |  |  |  |  |  |
|  | Calibration |  |  |  |  |  |  |
|  | Software licenses |  |  |  |  |  |  |
|  | Support contract |  |  |  |  |  |  |
| TOTAL |  |  |  |  |  |  |  |

Notes:

- Include all costs in [USD or local currency]

- Assume normal operation, no catastrophic failures

- Include inflation if multi-year contract

- Separate optional items clearly

## Pacific Requirements

All upper air systems must meet these additional requirements:

| Requirement | Specification | Compliant? |
| --- | --- | --- |
| Corrosion protection | Marine-grade materials, sealed connectors, tropical coating |  |
| UV resistance | All outdoor components UV-stabilized |  |
| Cyclone resistance | Antenna mast rated to 150 mph (240 km/h) winds |  |
| Lightning protection | Complete system protection with redundancy |  |
| Power autonomy | Minimum 8 hours backup for ground system |  |
| Temperature range | Operational: 0°C to +50°C (indoor), -10°C to +55°C (outdoor) |  |
| Humidity tolerance | 0-100% condensing environment |  |
| Salt spray protection | All outdoor components sealed to IP65 minimum |  |
| Mould/fungus resistance | All materials treated or selected for tropical resistance |  |
| Insect/pest protection | Screened ventilation, sealed cable entries |  |
| Gas safety | Ventilation, detection, PPE for tropical enclosed spaces |  |

## Compliance Checklist

Before finalising your tender, verify:

| Check | Completed? |
| --- | --- |
| Tier selection matches your operational needs | ☐ |
| System architecture fully defined | ☐ |
| All sensor specifications are included | ☐ |
| Data transmission requirements specified | ☐ |
| Power system defined | ☐ |
| Gas storage and safety addressed | ☐ |
| Physical infrastructure requirements clear | ☐ |
| Software and integration requirements specified | ☐ |
| Spare parts kit matches tier requirement | ☐ |
| Consumables quantified | ☐ |
| Field verification equipment included | ☐ |
| Warranty and SLA requirements specified | ☐ |
| Training requirements fully defined and specified | ☐ |
| Technical documentation requirements included | ☐ |
| Pacific requirements included | ☐ |
| Lifecycle cost worksheet completed | ☐ |
| Budget confirmed for 5-year operation | ☐ |

## Additional Information for Bidders

  - Site Information (to be completed by NMHS)

| Site | Latitude | Longitude | Elevation | Access | Power | Comms | Notes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 |  |  |  |  |  |  |  |
| 2 |  |  |  |  |  |  |  |
| 3 |  |  |  |  |  |  |  |

  - Existing Equipment at Sites

| Site | Existing Logger | Existing Sensors | Compatible? |
| --- | --- | --- | --- |
| 1 |  |  |  |
| 2 |  |  |  |

  - Special Local Requirements

Import duties: ☐ NMHS responsible ☐ Bidder responsible

Customs clearance: ☐ NMHS assists ☐ Bidder responsible

Local taxes: ☐ Applicable ☐ Exempt (specify)

Work permits: ☐ Required ☐ Not required

Frequency licenses: ☐ NMHS provides ☐ Bidder obtains ☐ Not required

Gas import permits: ☐ Required ☐ Not required

Site access restrictions: _________________________________